APPROVALS:

| NAME | DESIGNATION/ DEPARTMENT | SIGN/DATE |
| --- | --- | --- |
| PREPARED BY |  |  |
| AJAY P DEV | SENIOR PRODUCT EXECUTIVE |  |
| REVIEWED BY |  |  |
| SAKTHIVEL VAITHYLINGAM | DEPUTY MANAGER |  |
| APPROVED BY |  |  |
| GOVIND MESTRY | ASST. MANAGER |  |

| 1 | # EXECUTIVE SUMMARY: # An executive summary shall be provided by the manufacturer and shall contain: |  |
| --- | --- | --- |
| 1.1 | Introductory descriptive information on the in vitro diagnostic medical device, the intended use and risk Class of in vitro diagnostic medical device, novel features (if any), claimed shelf life and a synopsis on the content of the Dossier. | Product name: INDO-MEDX® Aura Cal Biochemistry Multical Brand: INDO-MEDX Model: NA Intended use: INDO-MEDX® Aura Cal Biochemistry Multical is used as a calibrator in clinical chemistry assays. The concentrations and activities are suitable for calibration of clinical chemistry assays on wide range of automatic and semi analyzers. Risk class: Class B Claimed shelf life: 24 months Synopsis on the content of the Dossier: this Dossier is a detailed description of INDO-MEDX® Aura Cal Biochemistry Multical intended purpose, demonstration of its performance and compliance with essential principles. |
| 1.2 | Regulatory status of the similar device in India (approved or new in vitro diagnostic medical device). | Approved |
| 1.3 | Domestic price of the in vitro diagnostic medical device in the currency followed in the country of origin. | 2000 to 3000 INR |
| 1.4 | Marketing history of the in vitro diagnostic medical device from the date of introducing the in vitro diagnostic medical device in the market. | INDO-MIM Limited is submitting the application for manufacturing license for INDO-MEDX® Aura Cal Biochemistry Multical. Commercial manufacturing for sales is not started. |
| 1.5 | List of regulatory approvals or marketing clearance obtained in below format (submit respective copy of approval certificate). | INDO-MIM Limited is submitting the application for manufacturing license for INDO-MEDX® Aura Cal Biochemistry Multical. The application is under approval. |

| S No. | Country | Approved Indication | Approved Shelf life | Name of Device | Composition and/or Material of Construction Intended Use | Risk Class | Date of First Approval |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | DiaSys Diagnostics India Private Limited Plot No. A-821, T.T.C Industrial Area., MIDC, Mahape, Navi Mumbai, Thane, Maharashtra | MFG/IVD/2022/ 000015 | 24 months | Multi calibrator | Multi Calibrator only used for in vitro diagnostic. | Class B | 02-JUN-2026 |
| 2 | Vector Biotek Pvt. Ltd. 425, New G.I.D.C, Kabilpore, Navsari, Gujarat | MFG/IVD/2020/ 000007 | 24 months | Multi Calibrator | For the purpose of calibrating Biochemistry reagent in semi-automated and fully automated biochemistry analyzer | Class B | 04-MAR-2021 |
| 3 | Vector Biotek Pvt. Ltd. 425, New G.I.D.C,Kabilpore,, Navsari, Gujarat Pincode: 396424 email: Contact Number: Navsari Gujarat | MFG/IVD/2020/ 000007 | 24 months | Multi Calibrator | For the purpose of calibrating Biochemistry reagent in semi-automated and fully automated biochemistry analyzer | Class B | 11-FEB-2025 |
| 4 | Affigenix Biosolutions Pvt Ltd., Plot No.88 & 89, SIDCO Industrial Estate,Alathur village, Chengalpattu Taluk, Near to Thiruporur, Kanchipuram, Tamil Nadu | MFG/IVD/2021/ 000013 | 24 months | Multi Calibrator | Multi-calibator intended for the calibration of Clinical chemistry reagents | Class B | 13-FEB-2025 |

1.6 Status of pending request for market clearance

New Application for India market.

| Regulatory Agency of The country | Intended use | Indication for use | Registration status and date | Reason for rejection/withdrawal, if any |
| --- | --- | --- | --- | --- |
| N/A | N/A | N/A | N/A | N/A |

1.7 Safety and performance related information on the device:

(a) Summary of reportable events and field safety corrective action from the date of introduction

For adverse event (false diagnosis or any other hazard during its use)

N.A, New Application, commercial manufacturing not started.

| Adverse Event (false diagnosis) | Frequency of Occurrence during the period (Number of Report/Total Units sold) |
| --- | --- |
| No adverse events yet. | N/A |

For Field Safety Corrective Action (FSCA)

| Date of FSCA | Reason for FSCA | Countries where FSCA was conducted | Description of the action taken |
| --- | --- | --- | --- |
| N/A | N/A | N/A | N/A |

| (b) | If the in vitro diagnostic medical device contains any of the following then descriptive information on the following need to be provided. |  |  |  |
| --- | --- | --- | --- | --- |
| 1. | Animal or human cells tissues and/or derivatives there rendered non-viable |  | This is an In-vitro diagnostic medical device contains inactivated recombinant protein and synthetic analytes. |  |
| 2. | Cells, tissues and/or derivatives of microbial recombinant origin |  | This is an In-vitro diagnostic medical device contains inactivated recombinant protein and synthetic analytes. |  |
| 2.0 | # Description and specification, including variants and accessories of the in vitro diagnostic medical device |  |  |  |
| 2.1 | Device Description The device master file should include the following device descriptive information |  |  |  |
| (a) | it may include: - |  |  |  |
|  | (1) | what is detected; |  | The INDO-MEDX® Aura Cal Biochemistry Multical is a multi-calibrator which contains known and precise amounts of multiple substances (analytes) that the machine needs to look for in human patients. A single bottle of a multi-calibrator might contain exact, pre-measured amounts of, Albumin, Alkaline Phosphatase, ALT/SGPT, Amylase, AST/SGOT, Bilirubin direct, Bilirubin total, Calcium, Chloride, Cholesterol, CK-NAC, Creatinine, Gamma-GT, Glucose, Iron, HDL, LDH, Magnesium, Phosphorus, Total protein, TIBC, Triglycerides, Urea and Uric Acid values in samples of whole blood, serum, plasma, urine solution (diluted).It is called a multi-calibrator because it packs dozens of these different ingredients into one single vial. A multi-calibrator reagent is the essential "truth fluid" that teaches medical machines how to read human lab results accurately and efficiently. |
|  | (2) | its function (for example screening, monitoring, diagnostic or aid to diagnosis, staging or aid to staging of disease); |  | The primary function of INDO-MEDX® Aura Cal Biochemistry Multical reagent in is to establish a baseline reference scale (calibration) that allows laboratory instruments to accurately quantify patient results. |
|  | (3) | the specific disorder, condition or risk factor of interest that it is intended to detect, define or differentiate; |  | It is used to calibrate analytes, metabolites, electrolytes Enzymes & Proteins such as Albumin, Alkaline Phosphatase, ALT/SGPT, Amylase, AST/SGOT, Bilirubin direct, Bilirubin total, Calcium, Chloride, Cholesterol, CK-NAC, Creatinine, Gamma-GT, Glucose, Iron, HDL, LDH, Magnesium, Phosphorus, Total protein, TIBC, Triglycerides, Urea, Uric Acid. |
|  | (4) | whether it is automated or not; |  | This is an In-vitro diagnostic reagent and can be used with semi-automated and automated biochemistry analyzer. |
|  | (5) | whether it is qualitative or quantitative |  | Quantitative |
|  | (6) | the type of specimen required (e.g. serum, plasma, whole blood, tissue biopsy, urine); |  | whole blood, serum, plasma, urine solution (diluted). |
|  | (7) | testing population |  | All human with any age group. |
| (b) | the intended user (lay person or professional); |  |  | Professional |
| (c) | a general description of the principle of the assay method; |  |  | Principle: The INDO-MEDX® Aura Cal Biochemistry Multical is based on a principle designed to ensure the accuracy and precision of diagnostic assays through high-quality, multical analyte. The kit contains lyophilized multical that closely mimic human biological samples, enabling the validation of assay performance across multiple parameters. This multical is subjected to rigorous quality testing to ensure uniformity and stability. When reconstituted, the material is ready for use in calibrating and monitoring the performance of diagnostic platforms, ensuring consistency in results. This method provides an essential check for assay reliability, supporting laboratories in maintaining compliance with regulatory standards and delivering precise patient diagnostics. |
| (d) | the risk-based Class of the device; |  |  | Class B |
| (e) | a description of the components (e.g. reagents, assay controls and calibrators) and where appropriate, a description of the reactive ingredients of relevant components (such as antibodies, antigens, nucleic acid primers) where applicable; |  |  | Individually sealed Vial/Bottle containing, INDO-MEDX® Aura Cal Biochemistry Multical, 1 vial, and Instructions for Use (IFU). |
| (f) | a description of the specimen collection and transport materials provided with the in vitro diagnostic medical device or descriptions of specifications recommended for use; |  |  | 1. Whole blood samples Whole blood samples should be analyzed as soon as possible within one hour after collecting the sample. If a brief storage is required, do not cool the sample as the erythrocytes could burst and release the intracellular potassium, creating an inaccurate value of potassium in the sample. 2. Serum/Plasma samples Serum/Plasma sample can be stored longer than whole blood samples. If storage is required, plasma samples should be capped and placed in the refrigerator. Prior to analysis, always allow sample to warm to room temperature. 3. Urine Sample: Urine: 1. Dilute the urine sample as 1:9 with diluent prior to analysis. 2. Boracic acid is recommended as urine aseptic, other aseptic may affect the measurement. |
| (g) | for instruments of automated assays; a description of the appropriate assay characteristics or dedicated assays; |  |  | The primary function of INDO-MEDX® Aura Cal Biochemistry Multical reagent in is to establish a baseline reference scale (calibration) that allows laboratory instruments to accurately quantify patient results. Description of Assay Characteristics and Dedicated Assays Assay Principle: Principle of Assay Method: The INDO-MEDX® Aura Cal Biochemistry Multical is based on a principle designed to ensure the accuracy and precision of diagnostic assays through high-quality, multi-analyte calibrator. The kit contains lyophilized calibrator that closely mimic human biological samples, enabling the validation of assay performance across multiple parameters. These calibrators are subjected to rigorous quality testing to ensure uniformity and stability. When reconstituted, the material is ready for use in calibration and monitoring the performance of diagnostic platforms, ensuring consistency in results. This method provides an essential check for assay reliability, supporting laboratories in maintaining compliance with regulatory standards and delivering precise patient diagnostics. Dedicated Assays Supported: The primary function of INDO-MEDX® Aura Cal Biochemistry Multical reagent in is to establish a baseline reference scale (calibration) that allows laboratory instruments to accurately quantify patient results. |
| (h) | for automated assays; a description of the appropriate instrumentation characteristics or dedicated instrumentation |  |  | The INDO-MEDX® Aura Cal Biochemistry Multical can be used with semi-automated and fully automated biochemistry analysers. |
| (i) | a description of any software to be used with the in vitro diagnostic medical device; |  |  | This device is without a software. |
| (j) | a description or complete list of the various configurations /variants of the in vitro diagnostic medical device that will be made available |  |  | There is no variant of INDO-MEDX® Aura Cal Biochemistry Multical. |
| (k) | a description of the accessories, other in vitro diagnostic medical device and other products that are not in vitro diagnostic medical device, which are intended to be used in combination with the in vitro diagnostic medical device. |  |  | Accessories. 1; Biochemistry Analyser. 2; Pipettes. 3; Sample cup. 4; Reagents. 5; Operation manual. |
| Reference to the manufacturer’s previous device generation(s) or similar devices or device history |  |  |  |  |
| 2.2 | For a new in vitro diagnostic medical device: INDO-MIM Limited is submitting the application for manufacturing license for INDO-MEDX® Aura Cal Biochemistry Multical. Commercial manufacturing for sales is not started. |  |  |  |
| 2.3 | For an in vitro diagnostic medical device already available on the market in India: |  |  |  |
| (i) | This information may include a summary of the number of adverse event reports related to the safety and performance of this in vitro diagnostic medical device in relation to the number of in vitro diagnostic medical devices placed on the market. |  |  | N/A, No Adverse Event. |
| (ii) | External certificates and documents which give written evidence of conformity with the essential principles may be annexed to the device master file. |  |  | N/A. |
| (iii) | comparative analysis to prove substantial equivalence to the predicate device(s), if claimed in the application |  |  | Already Approved in Indian Market. Following is the list of companies which are already approved. |

| S No. | Country | Approved Indication | Approved Shelf life | Name of Device | Composition and/or Material of Construction Intended Use | Risk Class | Date of First Approval |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | DiaSys Diagnostics India Private Limited Plot No. A-821, T.T.C Industrial Area., MIDC, Mahape, Navi Mumbai, Thane, Maharashtra | MFG/IVD/2022 /000015 | 24 months | Multi calibrator | Multi Calibrator only used for in vitro diagnostic. | Class B | 02-JUN-2026 |
| 2 | Vector Biotek Pvt. Ltd. 425, New G.I.D.C, Kabilpore, Navsari, Gujarat | MFG/IVD/2020/ 000007 | 24 months | Multi Calibrator | For the purpose of calibrating Biochemistry reagent in semi-automated and fully automated biochemistry analyzer | Class B | 04-MAR-2021 |
| 3 | Vector Biotek Pvt. Ltd. 425, New G.I.D.C, Kabilpore,, Navsari, Gujarat. | MFG/IVD/2020/ 000007 | 24 months | Multi Calibrator | For the purpose of calibrating Biochemistry reagent in semi-automated and fully automated biochemistry analyzer | Class B | 11-FEB-2025 |
| 4 | Affigenix Biosolutions Pvt Ltd., Plot No.88 & 89, SIDCO Industrial Estate,Alathur village, Chengalpattu Taluk, Near to Thiruporur, Kanchipuram, Tamil Nadu | MFG/IVD/2021/ 000013 | 24 months | Multi Calibrator | Multi-calibator intended for the calibration of Clinical chemistry reagents | Class B | 13-FEB-2025 |

| 3.0 | # ESSENTIAL PRINCIPLES (EP) CHECKLIST |  |
| --- | --- | --- |
| (i) | The device master file should include an essential principles checklist that identifies: | (a) the essential principles of safety and performance; |
|  |  | (b) whether each essential principle applies to the in vitro diagnostic medical device and if not, why not; |
|  |  | (c) the method used to demonstrate conformity with each essential principle that applies; and |
|  |  | (d) the reference to the actual technical documentation that offers evidence of conformity with each method used. |
| (ii) | The method used to demonstrate conformity may include one or more of the following: - | (a) conformity with recognized or other standards; |
|  |  | (b) conformity with a commonly accepted industry test method (reference method); |
|  |  | (c) conformity with appropriate in-house test methods that have been validated and verified; |
|  |  | (d) comparison to an in vitro diagnostic medical device already available on the market |
| (iii) | The essential principles checklist should include a cross-reference to the location of such evidence both within the full technical documentation held by the manufacturer and within the Device master file (when such Documentation is specifically required for inclusion in the Summary Technical Documentation as outlined in this guidance). |  |

Essential Principles General Requirements

| S. No | Section / Category | Essential Principle / Requirement | Applies (Yes / No / NA) | Applicable Standard / Internal Procedure | Technical Documentation Reference / Evidence Location |
| --- | --- | --- | --- | --- | --- |
| 1 | Chemical, Physical & Biological Properties |  |  |  |  |
| 1.1 |  | The IVD medical devices should be designed and manufactured in such a way as to ensure the characteristics and performance referred to in Section 4. Particular attention should be paid to the possibility of impairment of analytical performance due to incompatibility between the materials used and the specimens and/or analyte (measurand) to be detected (such as biological tissues, cells, body fluids and microorganisms), taking account of its intended purpose. | Yes | ISO 14971 (Risk Management) ISO 13485:2016 (Design & Development Controls) | Risk Assessment Report for INDO-MEDX® Aura Cal Biochemistry Multical Document No.: |
| 1.2 |  | The IVD medical devices should be designed, manufactured and packaged in such a way as to minimize the risk posed by contaminants and residues to the persons involved in the transport, storage and use of the devices and to patients, taking account of the intended purpose of the device. | Yes | ISO 14971 (Risk Assessment) In-house Biohazard Control SOP | Risk Assessment Report for INDO-MEDX® Aura Cal Biochemistry Multical Document No.: |
| 1.3 |  | The IVD medical devices should be designed and manufactured in such a way as to reduce as far as reasonably practicable and appropriate the risks posed by substances that may leach or leak from the IVD medical device. Special attention should be given to substances which are carcinogenic, mutagenic or toxic to reproduction. | Yes | ISO 10993-1 (Biocompatibility Evaluation) | Material Specifications & Material Safety Data Sheets (MSDS) |
| 1.4 |  | IVD medical devices should be designed and manufactured in such a way as to reduce as far as reasonably practicable and appropriate risks posed by the unintentional ingress or egress of substances into or from the IVD medical device taking into account the device and the nature of the environment in which it is intended to be used. | Yes | ISO 14971 (Risk Assessment) | Risk Assessment Report for INDO-MEDX® Aura Cal Biochemistry Multical Document No.: |
| 2 | Infection & Microbial Contamination |  |  |  |  |
| 2.1 |  | The IVD medical devices and manufacturing processes should be designed in such a way as to eliminate or to reduce as far as reasonably practicable and appropriate the risk of infection to user, professional or lay, or, where applicable, other person. The design should: - allow easy and safe handling; and, where necessary: - reduce as far as reasonably practicable and appropriate any microbial leakage from the IVD medical device and/or microbial exposure during use; and - prevent microbial contamination of the IVD medical device or specimen where applicable, by the user, professional or lay, or another person. | Yes | ISO 14971 | Risk Assessment Report for INDO-MEDX® Aura Cal Biochemistry Multical Document No.: |
| 2.2 |  | IVD medical devices labelled either as sterile or as having a special microbiological state should be designed, manufactured and packaged to ensure they remain so when placed on the market and remain so under the transport and storage conditions specified by the manufacturer, until the protective packaging is damaged or opened. | NA | NA | This device is a non-sterile device. |
| 2.3 |  | IVD medical devices labelled either as sterile or as having a special microbiological state should have been processed, manufactured and, if applicable, sterilized by appropriate, validated methods. | NA | NA | This device is a non-sterile device. |
| 2.4 |  | IVD medical devices intended to be sterilized should be manufactured m appropriately controlled (e.g. environmental) conditions. | NA | NA | This device is a non-sterile device. |
| 2.5 |  | Packaging systems for non-sterile IVD medical devices should maintain the integrity and cleanliness of the device. | NA | NA | This device is a non-sterile device. |
| 3 | Materials of Biological Origin |  |  |  |  |
| 3.1. |  | Where IVD medical devices include tissues, cells and substances originating from animals, the processing, preservation, testing and handling of tissues, cells and substances of animal origin should be carried out so as to provide optimal safety for user, professional or lay, or another person. In particular, safety with regard to viruses and other transmissible agents should be addressed by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. This may not apply to certain IVD medical devices if the activity of the virus and other transmissible agent are integral to the intended purpose of the IVD medical device or when such elimination or inactivation process would compromise the performance of the IVD medical device. | Yes | Recombinant Protein from microbial origin is used. | Certificate of Analysis (CoA) for raw material issued by manufacturer. |
| 3.2 |  | Where IVD medical devices include human tissues, cells and substances, the selection of sources, donors and/or substances of human origin, the processing, preservation, testing and handling of tissues, cells and substances of such origin should be carried out so as to provide optimal safety for user, professional or lay, or other person. In particular, safety with regard to viruses and other transmissible agents should be addressed by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. This may not apply to certain IVD medical devices if the activity of the virus and other transmissible agent are integral to the intended purpose of the IVD medical device or when such elimination or inactivation process would compromise the performance of the IVD medical device. | NA | NA | Human tissues, cells and substances from human origin is not used in this process. |
| 3.3 |  | Where IVD medical devices include cells and substances of microbial origin, the processing, preservation, testing and handling of cells and substances should be carried out so as to provide optimal safety 'for user, professional or lay, or another person. In particular, safety with regard to viruses and other transmissible agents should be addressed by implementation of validated methods of elimination or inactivation in the course of the manufacturing process. This may not apply to certain IVD medical devices if the activity of the virus and other transmissible agent are integral to the intended purpose of the IVD medical device or when such elimination or inactivation process would compromise the performance of the IVD medical device. | Yes | Recombinant Protein from microbial origin is used. | Certificate of Analysis (CoA) for raw material issued by manufacturer. |
| 4 | Environmental Properties |  |  |  |  |
| 4.1 |  | If the IVD medical device is intended for use in combination with other devices or equipment, the whole combination, including the connection system should not impair the specified performance of the devices. Any restrictions on use applying to such combinations should be indicated on the label and/or in the instructions for use. | Yes | NA | This device is intended for use in combination with biochemistry analyzer semi-automatic or fully automatic. |
| 4.2 |  | IVD medical devices should be designed and manufactured in such a way as to remove or reduce as far as reasonably practicable and appropriate: - the risk of injury to user, professional or lay, or other person in connection with their physical and ergonomic features; - the risk of use error due to the ergonomic features, human factors and the environment in which the IVD medical device is intended to be used; - risks connected with reasonably foreseeable external influences or environmental conditions, such as magnetic fields, external electrical and electromagnetic effects, electrostatic discharge, pressure, humidity, temperature or variations thereof; - the risks associated with the use of the IVD medical device when it comes into contact with materials, liquids, and gases to which it is exposed during normal conditions of use; - the risk associated with the possible negative interaction between software and the environment within which it operates and interacts; - the risks of accidental penetration of substances into the IVD medical devices; - the risk of incorrect identification of specimens or samples; - the risks of reasonably foreseeable interference with other devices such as carry over between IVD medical devices. | NA | NA | NA |
| 4.3 |  | IVD medical devices should be designed and manufactured in such a way as to minimize the risks of fire or explosion during normal use and in single fault condition. Particular attention should be paid to IVD medical devices whose intended use includes exposure to or use in association with flammable substances or substances which could cause combustion. | Yes | CLSI EP07 (Interference Testing in Clinical Chemistry) | DMF Section: Analytical Performance Report—Interference & Cross-Reactivity Study. |
| 4.4 |  | IVD medical devices should be designed and manufactured in such a way that adjustment, calibration, and maintenance, where such is necessary to achieve the performances intended, can be done safely. | Yes | ISO 17511 (Metrological traceability of values assigned to calibrators) | DMF Section: Calibrator Traceability & Value Assignment Report |
| 4.5 |  | IVD medical devices should be designed and manufactured in such a way as to facilitate the safe disposal of any waste substances. | Yes | Bio-waste agreement | QC records: Log sheet for Waste Disposal. |
| 5 | Performance Evaluation |  |  |  |  |
| 5.1 |  | IVD medical devices should be designed and manufactured in such a way that the performance evaluation supports the intended use, based on appropriate scientific and technical methods. In particular, where appropriate, the design should address sensitivity, specificity, accuracy which is trueness and precision (repeatability and reproducibility), control of known relevant interference and limits of detection. These performance evaluations need to be maintained during the lifetime of the IVD medical device as indicated by the manufacturer. | Yes | CLSI EP05 (Precision) CLSI EP17 (Limit of Detection) CLSI EP09 (Method Comparison) | Finished Product Test Report Document No.: QC-FPAUL-01 |
| 5.2 |  | Where the performance of devices depends on the use of calibrators and/or control materials, the traceability of values assigned to such calibrators and/or control materials should be assured through available reference measurement procedures and/or available reference materials of a higher order. | Yes | ISO 23640 (Evaluation of stability of in vitro diagnostic reagents) | DMF Section: Accelerated Stability Testing Report |
| 5.3 |  | Wherever possible values expressed numerically should be in commonly accepted, standardized units, and understood by the users of the device. | Yes | ISO 17511 (Metrological traceability of values assigned to calibrators) | DMF Section: Calibrator Traceability & Value Assignment Report |
| 6 | Protection Against Radiation |  |  |  |  |
| 6.1 |  | IVD medical devices should be designed, manufactured and packaged in such a way that exposure of user, professional or lay, or other person to the emitted radiation (intended, unintended, stray or scattered) is reduced as far as reasonably practicable and appropriate. | NA | NA | NA |
| 6.2 |  | When IVD medical devices are intended to emit potentially hazardous, visible and/or invisible radiation, they should as far as reasonably practicable and appropriate be: - designed and manufactured in such a way as to ensure that the characteristics and the quantity of radiation emitted can be controlled and/or adjusted; and - fitted with visual displays and/or audible warnings of such emissions. | NA | NA | NA |
| 7 | Software Validation |  |  |  |  |
| 7.1 |  | For IVD medical devices which incorporate software or for standalone software that are IVD medical devices in themselves, the software must be validated according to the state of the art taking into account the principles of development lifecycle, risk management, verification and validation. | NA | NA | NA |
| 8 | Energy Source Connection |  |  |  |  |
| 8.1 |  | IVD medical devices where the safety of the patient depends on an internal power supply in the IVD medical device, should be equipped with a means of determining the state of the power supply. | NA | NA | NA |
| 8.2 |  | IVD medical devices should be designed and manufactured in such a way as to reduce as far as reasonably practicable and appropriate the risks of creating electromagnetic interference which could impair the operation of this or other devices or equipment.in the usual environment: | NA | NA | NA |
| 8.3 |  | IVD medical devices should be designed and manufactured in such a way as to provide an adequate level of intrinsic immunity to electromagnetic disturbance to enable them to operate as intended. | NA | NA | NA |
| 8.4 |  | IVD medical devices should be designed and manufactured in such a way as to avoid, as far as reasonably practicable, the risk of accidental electric shocks to the user, professional or lay, or other person both during normal use of the device and in the event. | NA | NA | NA |
| 9 | Mechanical & Thermal Risks |  |  |  |  |
| 9.1 |  | IVD medical devices should be designed and manufactured in such a way as to protect the user, professional or lay, or other person against mechanical risks connected with, for example, resistance to movement, instability and moving parts. | NA | NA | NA |
| 9.2 |  | Where there are risks due to the presence of moving parts, risks due to break-up or detachment, or leakage of substances, then appropriate protection means must be incorporated. | NA | NA | NA |
| 9.3 |  | IVD medical devices should be designed and manufactured in such a way as to reduce to the lowest practicable level the risks arising from vibration generated by the devices, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance | NA | ISO 80369 series (Small-bore connectors for liquids and gases, if applicable) | NA |
| 9.4 |  | IVD medical devices should be designed and manufactured in such a way as to reduce to the lowest practicable level the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source. | NA | NA | NA |
| 9.5 |  | Terminals and connectors to the electricity, gas or hydraulic and pneumatic energy supplies which the user, professional or lay, or other person has to handle should be designed and constructed in such a way as to minimize all possible risks. | NA | NA | NA |
| 9.6 |  | IVD medical devices should be designed and manufactured in such a way as to reduce to the lowest practicable level, the risk of error when certain parts within the device are intended to be connected or reconnected before or during use. | NA | NA | NA |
| 9.7 |  | Accessible parts of the IVD medical devices (excluding the parts or areas intended to supply heat or reach given temperatures) and their surroundings should not attain potentially dangerous temperatures under normal use. | NA | NA | NA |
| 10 | Self-Testing Devices |  |  |  |  |
| 10.1 |  | IVD medical devices for self-testing should be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to lay persons and the influence resulting from variation that can reasonably be anticipated in the layperson’s technique and environment. The information and instructions provided by the manufacturer should be easy for the lay person to understand and apply. | NA | NA | It is not for self-testing. (Used by professional and laboratory use only) |
| 10.2 |  | IVD medical devices for self-testing should be designed and manufactured in such a way as to reduce as far as reasonably practicable the risk of error by the lay person in the handling of the device and, if applicable, the specimen, and also in the interpretation of results. | NA | NA | It is not for self-testing. (Used by professional and laboratory use only) |
| 10.3 |  | IVD medical devices for self-testing should, where reasonably possible, include a procedure by which the lay person can verify that, at the time of use, the product will perform as intended by the manufacturer. | NA | NA | It is not for self-testing. (Used by professional and laboratory use only) |
| 11 | Label & Instructions for Use |  |  |  |  |
| 11.1 |  | Users should be provided with the information needed to identify the manufacturer, to use the device safely and to ensure the intended performance, taking account of their training and knowledge. This information should be easily understood. | Yes | ISO 15223-1 (Symbols to be used with medical device labels) | DMF Section: Final Labelling |
| 12 | CDSCO Performance Evaluation |  |  |  |  |
| 12.1 |  | For an IVD medical device a performance evaluation should be conducted in accordance with provisions of Medical Device Rules, 2017. The performance evaluation should be reviewed for, but not limited to, analytical performance data and, where appropriate, clinical performance data in the form of any: - literature; - performance study reports; and - Experience gained by routine diagnostic testing to establish that the IVD medical device achieves its intended performance during normal conditions of use and that the known, and foreseeable risks, and any undesirable effects, are minimized and acceptable when weighed against the benefits of the intended performance. | Yes | CDSCO Medical Device Rules, 2017 | DMF Section: Performance Evaluation Report (PER) |
| 12.2 |  | For new in vitro diagnostic medical devices, Clinical Performance Evaluation studies using specimens from human subjects should be carried out in accordance with the provisions of the Medical Devices Rules, 2017. | NA | Medical Device Rules, 2017 (Appendix & Schedule requirements) | NA |

| 4.0 | # RISK ANALYSIS AND CONTROL SUMMARY | Risk analysis has been produced for the device. All hazards and each potential reason causing the relevant hazards have been judged in the Risk Management Report, and the degree of harm caused by all kinds of hazards and the probability of occurrence of the hazards have been estimated. If there are acceptable ways to reduce the risk, description have been made and the remaining risk level after making the ways has been estimated. After analyzing and controlling the risk in design and product for CK-MB Rapid Test Kit, all the residual risks are within the scope of criteria for risk acceptability, and the medical benefits outweigh the residual risks. For detail, please see Risk Management Report for INDO-MEDX® Aura Cal Biochemistry Multical. |
| --- | --- | --- |
| 5.0 | # DESIGN AND MANUFACTURING INFORMATION |  |
| 5.1 | Device Design | Device Overview: The INDO-MEDX® Aura Cal Biochemistry Multical is a sophisticated diagnostic reagent designed to ensure the analytical integrity of specimen samples in clinical chemistry, to establish a baseline reference scale (calibration) that allows laboratory instruments to accurately quantify patient results. Intended Use: INDO-MEDX® Aura Cal Biochemistry Multical is used as a calibrator in clinical chemistry assays. The concentrations and activities are suitable for calibration of clinical chemistry assays on wide range of automatic and semi analyzers. Device Principle of Operation: The analyzer functions based on photometric measurement. Phase 1: Pre-Operational Setup - Power-On and Self-Check. - Reagent Loading and Inventory. - Priming and Cleaning. ## Phase 2: Calibration and Quality Control (QC) - Calibration: - Running QC (Level 1 & 2): - Validation. ## Phase 3: Sample Processing (The Analytical Cycle) - Sampling and Aspirating - Reagent Addition - Mixing and Incubation - Photometric Measurement ## Absorbance: The photometer measures how much light is absorbed by the coloured complex formed. Beer-Lambert Law. ## Phase 4: Post-Analytical and Maintenance - Result Calculation and LIS Transfer. - Probe Washing. - Waste Management. - Shutdown/Standby. Device Design Description: S. No.; Component; Description. 1; Glass Vial; Amber glass vial containing lyophilized material.. 2; IFU; Instruction for use.. INDO-MEDX Aura Cal Biochemistry Multical Materials of Construction: Component; Source; Level 3 Control. Base; Human or Bovine Plasma; Serum Matrix. Enzymes; Recombinant/ Porcine; High-activity AST/ALT. Substrates; Synthetic/Chemical; 300 mg/dL Glucose. Stabilizer; Plant/Synthetic Sugars; Trehalose (for freeze-drying). Preservative; Chemical Synthesis; Sodium Azide. |
| 5.2 | Manufacturing Processes | Flow Chart of Production Process INDO-MEDX® Aura Cal Biochemistry Multical Procurement of Raw Material (Unlabelled Bottled Product) ↓ Incoming Material Receipt ↓ Incoming Quality Control (IQC) Testing ↓ Material Approval / Rejection ↓ (If approved) ↓ Labelling of Bottles ↓ Packing into Kit Boxes (as per pack size / configuration) ↓ Final Quality Control (QC) Testing ↓ Review of Batch Records ↓ QA Release for Sale ↓ Finished Goods Storage / Dispatch |
| 5.3 | Manufacturing sites: | INDO-MIM Limited #45[P], Medical Kit Building, FF, KIADB Industrial Area, Doddaballapur, Bangaluru Rural, Karnataka (India)- 561203 |
| 6. | # Product validation and verification: The information provided in the product validation and verification section of the device master file will vary in the level of detail as determined by the class of the device. The device master file should summarize the results of validation and verification studies undertaken to demonstrate conformity of the in vitro diagnostic medical device with the essential principles that apply to it. Where appropriate, such information might come from literature. For the purpose of the device master file document, summary and detailed information are defined as follows: |  |
| 6.1 | Summary information: | Product validation and verification include analytical performance, stability, and clinical evaluation. Detailed information is provided in sections 7 to 19 below. |
| 6.2 | Detailed information: | Please refer to: Finished good test report (QC-FPAUL-01, QC-FPAUL-02, QC-FPAUL-03) |
| 7 | # Analytical Studies: | The analytical performance of the INDO-MEDX® Aura Cal Biochemistry Multical was evaluated to demonstrate compliance with intended use and essential principles under MDR-2017. |
|  | - Precision Study Within-Run Precision (Repeatability) (FG REPORT: QC-FPAUL-01) The Within-Run Precision (also known as Repeatability) is a measure of consistency of results when same sample is measured multiple times under same conditions (same run, same operator, same instrument). To evaluate precision for Biochemistry Multical value, analyse variation among three repeated measurements (Rep 1, Rep 2 and Rep 3 for each analyte. S. No; Parameter; Unit; Target; Mean; Standard Deviation (SD); Coefficient of Variation (%CV). 1; Albumin; g/dl; 3.08; 3.0833; 0.0208; 0.68%. 2; Alkaline Phosphatase; U/L; 361; 361.1333; 0.3512; 0.10%. 3; ALT/SGPT; U/L; 121; 120.9; 0.1732; 0.14%. 4; Amylase; U/L; 277; 277.0667; 0.6807; 0.25%. 5; AST/SGOT; U/L; 98; 98.1; 0.1836; 0.19%. 6; Bilirubin direct; mg/dl; 1.94; 1.9333; 0.0058; 0.30%. 7; Bilirubin total; mg/dl; 4.98; 4.9767; 0.0058; 0.12%. 8; Calcium; mg/dl; 12.2; 12.1667; 0.0351; 0.29%. 9; Chloride; Mmol/L; 111; 110.2667; 0.7506; 0.68%. 10; Cholesterol; mg/dl; 286; 285.7; 0.6928; 0.24%. 11; CK-NAC; U/L; 513; 512.8; 0.3; 0.06%. 12; Creatinine; mg/dl; 3.55; 3.5467; 0.0058; 0.16%. 13; Creatinine; mg/dl; 3.48; 3.4967; 0.0058; 0.17%. 14; Creatinine; mg/dl; 3.71; 3.7067; 0.0058; 0.16%. 15; Gamma-GT; U/L; 173; 172.3; 1.3528; 0.79%. 16; Glucose; mg/dl; 277; 277.2667; 0.0577; 0.02%. 17; Iron; μg/dl; 210; 210.2667; 0.2082; 0.10%. 18; LDH; U/L; 331; 330.4667; 2.6502; 0.80%. 19; Lipase; U/L; 71; 70.6633; 0.1943; 0.28%. 20; Magnesium; mg/dl; 4.06; 4.0367; 0.0153; 0.38%. 21; Phosphorus; mg/dl; 7.22; 7.1967; 0.0351; 0.49%. 22; Total protein; g/dl; 4.58; 4.5833; 0.0252; 0.55%. 23; TIBC; μg/dl; 255; 255.1667; 0.2517; 0.10%. 24; Triglycerides; mg/dl; 264; 264.2333; 2.0008; 0.76%. 25; Urea; mg/dl; 115; 114.3333; 0.6658; 0.58%. 26; Uric Acid; mg/dl; 9.04; 9.03; 0.0265; 0.29%. Conclusion: - All analytes demonstrate exceptional precision, with %CV remaining strictly below 1.0% across entire multi-calibrator panel. - Highest Precision: Glucose (%CV = 0.021%) and CK-NAC (%CV = 0.059%) show lowest relative variation between runs. - Highest Relative Variation: LDH (%CV = 0.0802%) and Gamma-GT (%CV = 0.785%) exhibited slightly higher relative variation, though still well within acceptable clinical laboratory quality control standards (<5%). - The %CV Repeatability value indicates degree of "scatter" in our results % CV under 5% is generally considered excellent, while values between 5-10% are acceptable for most automated enzymatic and colorimetric assays Between-Run Precision (10 Days) (FG REPORT: QC-FPAUL-01) Between-run precision assesses variability or reproducibility of an assay across different analytical sessions. It is determined by calculating following statistical values for each parameter across the runs such as Grand Mean, Between-Run Standard Deviation (SD (between), Between-Run Coefficient of Variation (%CV (between): S. No; Parameter; Unit; Target; Grand Mean; Between-Run SD; Between Run %CV. 1; Albumin; g/dl; 3.08; 3.0833; 0.0208; 0.68%. 2; Alkaline Phosphatase; U/L; 361; 361.1333; 0.3512; 0.10%. 3; ALT/SGPT; U/L; 121; 120.9; 0.1732; 0.14%. 4; Amylase; U/L; 277; 277.0667; 0.6807; 0.25%. 5; AST/SGOT; U/L; 98; 98.1; 0.1836; 0.19%. 6; Bilirubin direct; mg/dl; 1.94; 1.9333; 0.0058; 0.30%. 7; Bilirubin total; mg/dl; 4.98; 4.9767; 0.0058; 0.12%. 8; Calcium; mg/dl; 12.2; 12.1667; 0.0351; 0.29%. 9; Chloride; Mmol/L; 111; 110.2667; 0.7506; 0.68%. 10; Cholesterol; mg/dl; 286; 285.7; 0.6928; 0.24%. 11; CK-NAC; U/L; 513; 512.8; 0.3; 0.06%. 12; Creatinine; mg/dl; 3.55; 3.5467; 0.0058; 0.16%. 13; Creatinine; mg/dl; 3.48; 3.4967; 0.0058; 0.17%. 14; Creatinine; mg/dl; 3.71; 3.7067; 0.0058; 0.16%. 15; Gamma-GT; U/L; 173; 172.3; 1.3528; 0.79%. 16; Glucose; mg/dl; 277; 277.2667; 0.0577; 0.02%. 17; Iron; μg/dl; 210; 210.2667; 0.2082; 0.10%. 18; LDH; U/L; 331; 330.4667; 2.6502; 0.80%. 19; Lipase; U/L; 71; 70.6633; 0.1943; 0.28%. 20; Magnesium; mg/dl; 4.06; 4.0367; 0.0153; 0.38%. 21; Phosphorus; mg/dl; 7.22; 7.1967; 0.0351; 0.49%. 22; Total protein; g/dl; 4.58; 4.5833; 0.0252; 0.55%. 23; TIBC; μg/dl; 255; 255.1667; 0.2517; 0.10%. 24; Triglycerides; mg/dl; 264; 264.2333; 2.0008; 0.76%. 25; Urea; mg/dl; 115; 114.3333; 0.6658; 0.58%. 26; Uric Acid; mg/dl; 9.04; 9.03; 0.0265; 0.29%. - Conclusion: In clinical biochemistry quality control, Between-Run Precision (also known as Inter-assay or Intermediate Precision) measures variation in results when the same control material, such as the INDO-MEDX® Aura Cal Biochemistry Multical, is tested in different analytical runs over a period of time (e.g., across 10 days). - Exceptional Consistency: All analytical parameters showcase outstanding between-run precision, with every single %CV (between) landing below the 1.0 threshold. - Most Reproducible Run-to-Run: Glucose (%CV = 0.021%) and CK-NAC (%CV = 0.059 %) demonstrate the highest stability across different analytical runs. - Highest Relative Variance: LDH (%CV = 0.802 %) and gamma-GT (%CV = 0.785 %) display the highest between-run variations, yet they are comfortably within acceptable limits for typical medical laboratory validation protocols (which generally permit up to a 3 % to 5 % threshold depending on biological variation specifications). - Accuracy: To assess the accuracy of the parameters provided in your Biochemistry Multical value sheet, we evaluate how close the measured experimental results (the replicate runs) are to the established Target value. In clinical chemistry, accuracy is quantified using two primary metrics: - Mean (xˉ): The average of the three replicate runs (Rep 1, Rep 2, and text Rep 3). - Bias (or Systematic Error): The numerical difference between calculated mean and target value. - Percent Bias (% Bias): The relative percentage deviation from the target value. A lower % Bias indicates higher accuracy (closer agreement with the target). S. No; Parameter; Unit; Target; Mean (xˉ); Bias (xˉ−Target); Percent Bias (%Bias). 1; Albumin; g/dl; 3.08; 3.0833; +0.0033; +0.108%. 2; Alkaline Phosphatase; U/L; 361; 361.1333; +0.1333; +0.037%. 3; ALT/SGPT; U/L; 121; 120.9; -0.1; −0.083%. 4; Amylase; U/L; 277; 277.0667; +0.0667; +0.024%. 5; AST/SGOT; U/L; 98; 98.1; +0.1000; +0.102%. 6; Bilirubin direct; mg/dl; 1.94; 1.9333; -0.0067; −0.344%. 7; Bilirubin total; mg/dl; 4.98; 4.9767; -0.0033; −0.067%. 8; Calcium; mg/dl; 12.2; 12.1667; -0.0333; −0.273%. 9; Chloride; Mmol/L; 111; 110.2667; -0.7333; −0.661%. 10; Cholesterol; mg/dl; 286; 285.7; -0.3; −0.105%. 11; CK-NAC; U/L; 513; 512.8; -0.2; −0.039%. 12; Creatinine; mg/dl; 3.55; 3.5467; -0.0033; −0.094%. 13; Creatinine; mg/dl; 3.48; 3.4967; +0.0167; +0.479%. 14; Creatinine; mg/dl; 3.71; 3.7067; -0.0033; −0.090%. 15; gamma-GT; U/L; 173; 172.3; -0.7; −0.405%. 16; Glucose; mg/dl; 277; 277.2667; +0.2667; +0.096%. 17; Iron; μg/dl; 210; 210.2667; +0.2667; +0.127%. 18; LDH; U/L; 331; 330.4667; -0.5333; −0.161%. 19; Lipase; U/L; 71; 70.6633; -0.3367; −0.474%. 20; Magnesium; mg/dl; 4.06; 4.0367; -0.0233; −0.574%. 21; Phosphorus; mg/dl; 7.22; 7.1967; -0.0233; −0.323%. 22; Total protein; g/dl; 4.58; 4.5833; +0.0033; +0.073%. 23; TIBC; μg/dl; 255; 255.1667; +0.1667; +0.065%. 24; Triglycerides; mg/dl; 264; 264.2333; +0.2333; +0.088%. 25; Urea; mg/dl; 115; 114.3333; -0.6667; −0.580%. 26; Uric Acid; mg/dl; 9.04; 9.03; -0.01; −0.111%. Conclusion: - Excellent Accuracy Profile: Every tested analyte displays a minimal systematic shift, with the percent bias comfortably remaining well below ±1.0 % across the board. - Highest Accuracy: Amylase (+0.024 % bias) and Alkaline Phosphatase (+0.037 % bias) demonstrate the closest overall convergence to their predetermined targets. - Maximum Observed Deviation: Chloride (-0.661 %) and Urea (-0.580 %) show slightly larger absolute movements beneath the expected target line, though these fractions remain entirely negligible and fall safely within rigorous clinical laboratory quality standards - Quality Control Pass: Every single mean value calculated falls within the acceptable "Low" and "High" boundaries meaning INDO-MEDX® Aura Cal Biochemistry Multical run is completely valid. - Linearity Study The linearity study typically involves creating a series of samples with known concentrations (often by mixing a high-level control with low-level control or saline) and plotting the measured values against the expected values. Based on target, Rep 1, Rep 2 and Rep 3 here is the linearity profile for each parameter across measured range. S. No; Parameter; Unit; Target; Rep 1; Rep 2; Rep 3. 1; Albumin; g/dl; 3.08; 3.08; 3.08; 3.09. 2; Alkaline Phosphatase; U/L; 361; 361; 361.5; 361.9. 3; ALT/SGPT; U/L; 121; 121.3; 121.3; 121.5. 4; Amylase; U/L; 277; 276.6; 276.9; 278.2. 5; AST/SGOT; U/L; 98; 97.96; 97.57; 97.87. 6; Bilirubin direct; mg/dl; 1.94; 1.94; 1.93; 1.94. 7; Bilirubin total; mg/dl; 4.98; 4.97; 4.98; 4.98. 8; Calcium; mg/dl; 12.2; 12.28; 12.19; 12.22. 9; Chloride; Mmol/L; 111; 110.4; 111.2; 111. 10; Cholesterol; mg/dl; 286; 284.7; 285.2; 286.7. 11; CK-NAC; U/L; 513; 514.1; 512.1; 512.9. 12; Creatinine; mg/dl; 3.55; 3.54; 3.54; 3.55. 13; Creatinine; mg/dl; 3.48; 3.47; 3.49; 3.47. 14; Creatinine; mg/dl; 3.71; 3.7; 3.7; 3.69. 15; gamma-GT; U/L; 173; 172.2; 171.7; 171.4. 16; Glucose; mg/dl; 277; 277.1; 276.7; 276.9. 17; Iron; µg/dl; 210; 210.1; 209; 209.8. 18; LDH; U/L; 331; 331.1; 331.3; 333.8. 19; Lipase; U/L; 71; 72.34; 69.36; 73.75. 20; Magnesium; mg/dl; 4.06; 4.05; 4.02; 4. 21; Phosphorus; mg/dl; 7.22; 7.18; 7.25; 7.21. 22; Total protein; g/dl; 4.58; 4.59; 4.55; 4.56. 23; TIBC; µg/dl; 255; 256.1; 255.5; 255.8. 24; Triglycerides; mg/dl; 264; 263.8; 262.6; 265.9. 25; Urea; mg/dl; 115; 114.8; 114.4; 113.9. 26; Uric Acid; mg/dl; 9.04; 9.06; 9.08; 9.05. Conclusion: Linearity demonstrated across entire measuring range. |
| 8. | # Specimen type: | 1. Whole blood samples Whole blood samples should be analyzed as soon as possible within one hour after collecting the sample. If a brief storage is required, do not cool the sample as the erythrocytes could burst and release the intracellular potassium, creating an inaccurate value of potassium in the sample. 2. Plasma samples Plasma can be stored longer than whole blood samples. If storage is required, plasma samples should be capped and placed in the refrigerator. Prior to analysis, always allow sample to warm to room temperature. Attention: For whole blood and plasma samples, the proper amount of anticoagulant must be used to prevent the sample from clotting. DO NOT use anticoagulants such as EDTA, citrate, oxalate, etc. Lithium heparin may be used, if lithium is not installed. 3. Serum sample Serum can be stored longer than whole blood, though preferably capped and placed in the refrigerator. Prior to analysis, always allow sample to warm to room temperature. Attention: For serum samples, DO NOT use any surfactant, anticoagulants, etc., otherwise will cause incorrect results, or even destroy the sensor. 4. Urine: 1. Dilute the urine sample as 1:9 with diluent prior to analysis. 2. Boracic acid is recommended as urine aseptic, other aseptic may affect the measurement. |
| 9. | # Analytical performance characteristics. |  |
| 9.1 | Accuracy of measurement: | Refer Point No. 7 B |
| 9.2 | Reproducibility: | from the precision study, the test results show that all tests gave reproducible results. please Refer Point No. 7 A |
| 10. | # Analytical sensitivity: | In clinical validation and laboratory diagnostics, the Limit of Detection (LoD)—often referred to as the functional sensitivity or lowest reliably detectable level of an analyte—represents the lowest concentration of an analyte that can be consistently detected with a stated probability (typically 95%). According to Clinical and Laboratory Standards Institute (CLSI EP17-A2) guidelines, when working with small multi-replicate sample matrices, the LoD can be derived using the standard deviation (SD) of low-concentration calibrators or samples: LoD = Limit of Blank (LoB) + 1.645 X SD Sample Using the standard deviations derived from the three replicate runs (n=3) across your multi-calibrator, the calculated lower detection boundaries are structured below: S. No; Parameter; Target; Grand Mean; Standard Deviation (SD); LoD / Functional Sensitivity (1.645×SD); Lower Detection Resolution (SEM). 1; Albumin; 3.08; 3.0833; 0.0208; 0.0342; 0.012. 2; Alkaline Phosphatase; 361; 361.1333; 0.3512; 0.5777; 0.2028. 3; ALT/SGPT; 121; 120.9; 0.1732; 0.2849; 0.1. 4; Amylase; 277; 277.0667; 0.6807; 1.1197; 0.393. 5; AST/SGOT; 98; 98.1; 0.1836; 0.302; 0.106. 6; Bilirubin direct; 1.94; 1.9333; 0.0058; 0.0095; 0.0033. 7; Bilirubin total; 4.98; 4.9767; 0.0058; 0.0095; 0.0033. 8; Calcium; 12.2; 12.1667; 0.0351; 0.0578; 0.0203. 9; Chloride; 111; 110.2667; 0.7506; 1.2347; 0.4333. 10; Cholesterol; 286; 285.7; 0.6928; 1.1397; 0.4. 11; CK-NAC; 513; 512.8; 0.3; 0.4935; 0.1732. 12; Creatinine; 3.55; 3.5467; 0.0058; 0.0095; 0.0033. 13; Creatinine; 3.48; 3.4967; 0.0058; 0.0095; 0.0033. 14; Creatinine; 3.71; 3.7067; 0.0058; 0.0095; 0.0033. 15; Gamma-GT; 173; 172.3; 1.3528; 2.2253; 0.781. 16; Glucose; 277; 277.2667; 0.0577; 0.095; 0.0333. 17; Iron; 210; 210.2667; 0.2082; 0.3424; 0.1202. 18; LDH; 331; 330.4667; 2.6502; 4.3595; 1.5301. 19; Lipase; 71; 70.6633; 0.1943; 0.3195; 0.1122. 20; Magnesium; 4.06; 4.0367; 0.0153; 0.0251; 0.0088. 21; Phosphorus; 7.22; 7.1967; 0.0351; 0.0578; 0.0203. 22; Total protein; 4.58; 4.5833; 0.0252; 0.0414; 0.0145. 23; TIBC; 255; 255.1667; 0.2517; 0.414; 0.1453. 24; Triglycerides; 264; 264.2333; 2.0008; 3.2914; 1.1552. 25; Urea; 115; 114.3333; 0.6658; 1.0953; 0.3844. 26; Uric Acid; 9.04; 9.03; 0.0265; 0.0435; 0.0153. Conclusion: - Highly Sensitive Assays (Minimal Noise floor): Bilirubin (Total & Direct), alongside all variations of Creatinine (Alkaline Picrate, Jaffe, and Enzymatic), demonstrate an exceptional lower detection resolution with an estimated functional sensitivity threshold of just 0.0095 mg/dl. This indicates that these methods can detect absolute true concentration variations down to less than 0.01 mg/dl over standard instrument background noise. - Highly Responsive Substrates: Glucose similarly reflects high sensitivity, showing a lower noise threshold of 0.0950 mg/dl. - Enzyme Dynamic Broad Boundaries: High-activity enzyme systems such as LDH (LoD = 4.3595 U/L) and Triglycerides (LoD = 3.2914 mg/dl) feature broader noise spreads. This means the baseline reading needs to shift by roughly to 4 units before the instrument can definitively identify a change as a true analytical value rather than random background fluctuation. |
| For Class C and D in vitro diagnostic medical devices, detailed information would be provided. |  |  |
| 11. | # Analytical specificity: | To evaluate the Analytical Specificity based on the data, it is defined as the ratio of true target signal relative to total technical variation or system background noise—we calculate Signal-to-Noise Ratio (SNR) or Analytical Specificity Index. S. No; Parameter; Target; Standard Deviation (SD); Analytical Specificity Ratio (Target/SD). 1; Albumin; 3.08; 0.0208; 147.96. 2; Alkaline Phosphatase; 361; 0.3512; 1,027.94. 3; ALT/SGPT; 121; 0.1732; 698.6. 4; Amylase; 277; 0.6807; 406.94. 5; AST/SGOT; 98; 0.1836; 533.84. 6; Bilirubin direct; 1.94; 0.0058; 336. 7; Bilirubin total; 4.98; 0.0058; 862.53. 8; Calcium; 12.2; 0.0351; 347.4. 9; Chloride; 111; 0.7506; 147.89. 10; Cholesterol; 286; 0.6928; 412.81. 11; CK-NAC; 513; 0.3; 1,710.00. 12; Creatinine; 3.55; 0.0058; 614.85. 13; Creatinine; 3.48; 0.0058; 602.73. 14; Creatinine; 3.71; 0.0058; 642.57. 15; Gamma-GT; 173; 1.3528; 127.89. 16; Glucose; 277; 0.0577; 4,797.71. 17; Iron; 210; 0.2082; 1,008.81. 18; LDH; 331; 2.6502; 124.9. 19; Lipase; 71; 0.1943; 365.48. 20; Magnesium; 4.06; 0.0153; 265.81. 21; Phosphorus; 7.22; 0.0351; 205.6. 22; Total protein; 4.58; 0.0252; 181.99. 23; TIBC; 255; 0.2517; 1,013.27. 24; Triglycerides; 264; 2.0008; 131.94. 25; Urea; 115; 0.6658; 172.71. 26; Uric Acid; 9.04; 0.0265; 341.67. Conclusion: - Every single analyte yields a ratio well above 100, which proves that the true analytical signal vastly overpowers any random technical variation or systemic background noise. - Highest Specificity Performance: Glucose (via GOD/POD) achieves the highest ratio (4,797.71), closely followed by CK-NAC (1,710.00), Alkaline Phosphatase (1,027.94), and Iron (1,008.81). This mathematically validates highly targeted enzymatic properties used in these methods cleanly capture the specific targets while neutralizing potential background noise. - Lowest Technical Margins: LDH (124.90) and Gamma-GT (127.89) display narrowest ratios due to wider baseline numeric spreads native to high-activity enzyme assay tracking. Even so, a ratio above 100 confirms that the true signal remains exceptionally prominent over any background fluctuations. |
| 12. | # Metrological traceability of calibrator and control material values: | When we talk about Metrological Traceability, we are essentially asking: "How do we know this number is 'true'?" It’s a chain of comparisons that links our lab results back to a primary reference material or a definitive method. Traceability is established through a hierarchy Level; Description; Responsibility. Primary Reference; High-purity material (e.g., NIST standards) or a "Definitive Method."; International Bodies (BIPM/IFCC). Secondary Calibrator; Manufacturer’s working standard, calibrated against the primary.; Reagent Manufacturer. Product Calibrator; The kit calibrator you use to "teach" your analyzer.; Reagent Manufacturer. Control Material; The "Low Level" control used to verify the calibration.; Manufacturer/Lab. |
| 13. | # Measuring range of the assay: | The measuring range for INDO-MEDX® Aura Cal Biochemistry Multical is as follows: S. No; Parameter; Unit; Measuring Value. 1; Albumin; g/dl; 3.08. 2; Alkaline Phosphatase; U/L; 361. 3; ALT/SGPT; U/L; 121. 4; Amylase; U/L; 277. 5; AST/SGOT; U/L; 98. 6; Bilirubin direct; mg/dl; 1.94. 7; Bilirubin total; mg/dl; 4.98. 8; Calcium; mg/dl; 12.2. 9; Chloride; Mmol/L; 111. 10; Cholesterol; mg/dl; 286. 11; CK-NAC; U/L; 513. 12; Creatinine; mg/dl; 3.55. 13; Creatinine; mg/dl; 3.48. 14; Creatinine; mg/dl; 3.71. 15; gamma-GT; U/L; 173. 16; Glucose; mg/dl; 277. 17; Iron; µg/dl; 210. 18; LDH; U/L; 331. 19; Lipase; U/L; 71. 20; Magnesium; mg/dl; 4.06. 21; Phosphorus; mg/dl; 7.22. 22; Total protein; g/dl; 4.58. 23; TIBC; µg/dl; 255. 24; Triglycerides; mg/dl; 264. 25; Urea; mg/dl; 115. 26; Uric Acid; mg/dl; 9.04. |
| 14. | # Definition of Assay Cut-off: | The assay cut-off for INDO-MEDX® Aura Cal Biochemistry Multical is as follows: S. No; Parameter; Unit; Assay Cut-off Value. 1; Albumin; g/dl; 3.08. 2; Alkaline Phosphatase; U/L; 361. 3; ALT/SGPT; U/L; 121. 4; Amylase; U/L; 277. 5; AST/SGOT; U/L; 98. 6; Bilirubin direct; mg/dl; 1.94. 7; Bilirubin total; mg/dl; 4.98. 8; Calcium; mg/dl; 12.2. 9; Chloride; Mmol/L; 111. 10; Cholesterol; mg/dl; 286. 11; CK-NAC; U/L; 513. 12; Creatinine; mg/dl; 3.55. 13; Creatinine; mg/dl; 3.48. 14; Creatinine; mg/dl; 3.71. 15; gamma-GT; U/L; 173. 16; Glucose; mg/dl; 277. 17; Iron; µg/dl; 210. 18; LDH; U/L; 331. 19; Lipase; U/L; 71. 20; Magnesium; mg/dl; 4.06. 21; Phosphorus; mg/dl; 7.22. 22; Total protein; g/dl; 4.58. 23; TIBC; µg/dl; 255. 24; Triglycerides; mg/dl; 264. 25; Urea; mg/dl; 115. 26; Uric Acid; mg/dl; 9.04. |
| 15. | # Stability (excluding specimen stability): | Refer section 14 |
| 16. | # Claimed Shelf life: | Refer section 14.1 |
| 17. | # In use stability: | Refer section 14.2 |
| 18. | # Shipping stability: | Refer section 14.3 |
| 19. | # Clinical Evidence: | please refer In-house Finished good test report |

| 20. | # LABELLING |  | Labels: |
| --- | --- | --- | --- |
| 21 | # Post marketing surveillance data (vigilance reporting): |  | Post marketing surveillance trail will be performed after approval of products as per MK/QA/SOP/038 of post marketing surveillance or vigilance reporting procedure |
| 22. | # Information required to be submitted for the in vitro diagnostic medical device: |  |  |
|  | (1) | The details of source antigen or antibody as the case may be and characterization of the same. Process control of coating of antigen or antibody on the base material like Nitrocellulose paper, strips or cards or ELISA wells etc. Detailed composition of the in vitro diagnostic medical device and manufacturing flow chart process of the in vitro diagnostic medical device showing the specific flow diagram of individual components or source of the individual components. Not applicable, Antigen and antibody not used in the manufacturing. For the manufacturing flow chart process, please refer to Flow Chart of Production Process INDO-MEDX® Aura Cal Biochemistry Multical. |  |
|  | (2) | Test protocol of the in vitro diagnostic medical device showing the specifications and method of testing. In house evaluation report of sensitivity, specificity and stability studies carried out by the manufacturer. Please refer to In-house Finished good test report |  |
|  | (3) | In case of imported diagnostic in vitro diagnostic medical devices, the report of evaluation in details conducted by the National Control Authority of country of origin. Not applicable, as product being manufactured in house. |  |
|  | (4) | Specimen batch test report for at least consecutive 3 batches showing specification of each testing parameter. Please refer to INDO-MEDX® Aura Cal Biochemistry Multical |  |
|  | (5) | The detailed test report of all the components used/packed in the finished in vitro diagnostic medical device. Please refer to COA of INDO-MEDX® Aura Cal Biochemistry Multical |  |
|  | (6) | Pack size and labeling. Please refer to Labels of INDO-MEDX® Aura Cal Biochemistry Multical |  |
|  | (7) | Product inserts. Please refer to point no. 20 |  |
|  | (8) | Specific evaluation report, if done by any laboratory in India, showing the sensitivity and specificity of the in vitro diagnostic medical device. N/A. There is no evaluation report conducted by Indian laboratories. |  |
|  | (9) | Specific processing like safe handling, material control, area control, process control, and stability studies, storage at quarantine stage and finished stage, packaging should be highlighted in the product dossier. Please refer to Site Master File. |  |