| IMDX-362 | - |
| --- | --- |
|  | 2-8°C |

For in vitro diagnostic use only.

## INTENDED USE

For use as a calibrator in clinical chemistry assays.

The concentrations and activities are suitable for calibration of clinical chemistry assays on wide range of automatic and semi analyzers.

## SAFETY PRECAUTIONS AND WARNINGS

Do not pipette by mouth. Exercise the normal precautions required for handling laboratory reagents.

Human source material, from which this product has been derived, has been tested at donor level for the Human Immunodeficiency Virus (HIV 1, HIV 2) antibody, Hepatitis B Surface Antigen (HbsAg), and Hepatitis C Virus (HCV) antibody and found to be NON REACTIVE. FDA approved methods have been used to conduct these tests.

However, since no method can offer complete assurance as to the absence of infectious agents, this material and all patient samples should be handled as though capable of transmitting infectious diseases and disposed of accordingly.

## STORAGE AND STABILITY

Unreconstituted serum is stable up to the expiry date shown on the side of each individual bottle. Once reconstituted, the components of the Calibration Sera are stable for 8 hours at

+15ºC to +25°C, 7 days at +2°C to +8°C, and 28 days at –20°C when frozen once (see limitations).

## LIMITATIONS

After reconstitution, Bicarbonate is stable for 8 hours in the closed bottle and 1 hour in the open bottle.

Alkaline Phosphatase is stable for 3 days at +2 ºC to +8°C and levels in the reconstituted serum will rise over the stability period. It is recommended that the reconstituted serum be allowed to stand for 1 hour at +15ºC to +25°C before measurement.

Bilirubin in the serum is light sensitive and it is recommended that the serum is stored in the dark. Stored in the dark, it is stable for1 day at +2°C to +8°C. Do not store at +15°C to +25°C. Do not freeze.

GLDH is stable for 1 day at +2ºC to +8°C.

AST (GOT) is stable for 2 days at +2°C to +8°C, 8 hours at +15°C to +25°C or 28 days when frozen once at -18°C to -24°C. Bacterial contamination of the reconstituted serum will cause reductions in the stability of many components. Different lot numbers of this calibrator should not be interchanged, as the values assigned to the calibrators vary from lot to lot.

Due to the zinc content in some batches of rubber stoppers,

the QC material should be aliquoted into suitable containers without rubber stoppers and stored at +2°C to +8°C to ensure stable zinc levels throughout the stability period.

## PREPARATION FOR USE

Serum must only be reconstituted using the following procedure:

- Open the vial carefully, avoiding any loss of material.

- Reconstitute by pipetting exactly 5 ml of distilled water at

+15ºC to +25ºC, into the vial.

- Replace the rubber stopper and leave to stand for 30 minutes out of bright light before use.

- Swirl gently several times during the reconstitution period to ensure that the contents are completely dissolved.

- Prior to use, mix the contents by inverting the vial. Do not

shake the vial as the formation of foam should be avoided. Ensure that no lyophilised material remains unreconstituted.

- The serum is then ready for use with either a manual test or with an automated instrument.

## MATERIALS PROVIDED

INDO-MEDX Aura Cal

## MATERIALS REQUIRED BUT NOT PROVIDED

Volumetric pipette, double deionized water.

## VALUE ASSIGNMENT

Each batch of serum is distributed to approximately 3000 laboratories worldwide and values are assigned by a consensus of results obtained by these laboratories. The Calibration values for each instrument have been determined in at least 10 independent laboratories. Values are verified against a master lot of calibrator, which is traceable to reference methods or reference materials. In some cases values may be assigned at INDO-MIM in comparison to a master lot of calibrator, which is traceable to reference methods or reference materials.

If an instrument specific value is not available, refer to the Method section. If necessary, contact INDO-MIM Technical Services, Tel: +91 8061418789 or

email caremedicalkits@indo-medx.com

# Value Sheet

| PARAMETER | METHOD | UNIT | VALUE |
| --- | --- | --- | --- |
| Albumin | Bromocresol Green | g/dl | 3.08 |
| Alkaline Phosphatase | AMP optimized to IFCC 37°C | U/L | 361 |
| ALT/SGPT | IFCC without PDP 37°C | U/L | 121 |
| Amylase | Direct (CNPG3) | U/L | 277 |
| AST/SGOT | IFCC without PDP 37°C | U/L | 98 |
| Bilirubin direct | Diazo Method | mg/dl | 1.94 |
| Bilirubin total | Diazo Method | mg/dl | 4.98 |
| Calcium | Arsenazo III Method | mg/dl | 12.2 |
| Chloride | Colorimetric | Mmol/L | 111 |
| Cholesterol | CHOD/POD Method | mg/dl | 286 |
| CK-NAC | IFCC | U/L | 513 |
| Creatinine | Alkaline Picrate | mg/dl | 3.55 |
| Creatinine | Jaffe | mg/dl | 3.48 |
| Creatinine | Enzymatic | mg/dl | 3.71 |
| Gamma-GT | IFCC | U/l | 173 |
| Glucose | GOD/POD Method | mg/dl | 277 |
| Iron | Colorimetric without ppt | µg/dl | 210 |
| LDH | L to P IFCC | U/l | 331 |
| Lipase | Colorimetric | U/l | 71 |
| Magnesium | Xylidyl Blue | mg/dl | 4.06 |
| Phosphorus | Phosphomolybdate UV | mg/dl | 7.22 |
| Total protein | Biuret End Point | g/dl | 4.58 |
| TIBC | Colorimetric | µg/dl | 255 |
| Triglycerides | GPO/PAP No Correction | mg/dl | 264 |
| Urea | GLDH Method | mg/dl | 115 |
| Uric Acid | AOX Method | mg/dl | 9.04 |