Device Master File For

In-Vitro Diagnostic Medical Device

Q-Line® Albumin

Table of Contents

1.0	Executive Summary.	4

2.0	Description and specification, including variants and accessories of the in vitro diagnostic medical device.	6

3.0	Essential Principles Checklist	9

4.0	Risk Analysis and Control Summary.	24

5.0	Design and Manufacturing Information	40

6.0	Product Validation and Verification	54

7.0	Analytical Studies	60

8.0	Specimen Type	60

9.0	Analytical Performance Characteristics	60

10.0	Analytical Sensitivity	61

11.0	Analytical Specificity	63

12.0	Metrological traceability of calibrator and control material values:	64

13.0	Definition of Assay Cut-off:	66

14.0	Stability (excluding specimen stability)	66

16.0	Claimed Shelf Life	66

17.0	In-Use Stability	68

18.0	Shipping Stability	68

19.0	Clinical Evidence	69

20.0	Labelling	70

21.0	Post Market Surveillance Data (Vigilance Reporting):	71

22.0	Information Required to be submitted for the in-vitro diagnostic medical device.	73

List of Tables

Table 6.1: Risk Analysis Product. Q-Line® Albumin	30

Table 6.2: Severity Table	30

Table 6.3: Probability Tables	30

Table 4 . Precision Table	56

Table 5 . Accuracy Table	56

Table 6: Analytical Sensitivity Study	62

Table 6: Analytical Specificity Study	63

Table 22.1. List Of Standards	76

List of Figures

Figure 6.1: Risk Assessment Matrix	24

Figure 6.2: Identify Hazards	25

Figure 6.3: Risk Analysis	26

Figure 6.4: Determining Risk Impact	26

Figure 6.5: Likelihood. the probability of a risk	27

Figure 6.6: Likelihood. the probability of a risk	28

Figure 6.7:Risk Estimation	31

Executive Summary.

An executive summary shall be provided by the manufacturer and shall contain:

  - Introductory descriptive information on the in vitro diagnostic medical device, the intended use and risk Class of in vitro diagnostic medical device, novel features (if any), claimed shelf life and a synopsis on the content of the dossier.

Intended Use: Q-Line® Albumin is intended for the quantitative in vitro diagnostic determination of albumin in human serum.

Class: It comes under Class B

Novel Feature: Albumin, synthetized primarily by the liver, represents 50 to 60% of total serum proteins. Because of its small size and its high plasmatic concentration, albumin is the major protein component of most extravascular body ﬂuid, including CSF, interstitial ﬂuid, urine and amniotic ﬂuid. Albumin's primary function is the maintenance of colloidal osmotic pressure in both extravascular and vascular spaces, with continuous equilibration. Albumin also binds and transports a large number of compounds (ions, free fatty acids, bilirubin, drugs...). Albumin is a mobile reserve of amino acids. Increased levels of albumin are present only in acute dehydration, especially critical for new born. Hypoalbuminemia is seen in a multitude of diseases bound to diﬀerent pathological states: (1) acute and chronic inﬂammation, (2) decreased synthesis: hepatic insuﬃciency, malnutrition, analbuminemia, (3) increased loss: nephritic syndrome, gastrointestinal loss, sever end large burns, bedsore, (4) increased catabolism: fever, hyperthyroidism etc.

Claimed Shelf life: 24 months.

  - Regulatory status of the similar device in India (approved or new in vitro diagnostic medical device).

The Regulatory status of the similar device in India is approved, this product is not a new in In-vitro diagnostic (IVD) medical device.

  - Domestic price of the in vitro diagnostic medical device in the currency followed in the country of origin.

The domestic price of this in vitro diagnostic medical device in the currency followed in the country of origin may range tentatively from ₹ 60-70 per kit.

  - Marketing history of the in vitro diagnostic medical device from the date of introducing the in vitro diagnostic medical device in the market.

This is not a new IVD product in market, it is being marketed by following other companies.

Siemens Healthcare Private Limited– Albumin

  - List of regulatory approvals or marketing clearance obtained in below format (submit respective copy of approval certificate).

| S. No. | Name of the country | Approved indication | Approved shelf life | Composition | Risk Class | Date of first approval |
| --- | --- | --- | --- | --- | --- | --- |
| 1 | India | Q-Line® Albumin | 24 months | Bromocresol Green (BCG), Buffer, Surfactants, Stabilizers, Product Insert | B |  |

  - Status of pending request for market clearance

No such request for market clearance is pending.

| Regulatory Agency of the country | Intended use | Indication For use | Registration status and date | Reason for rejection/ withdrawal if any |
| --- | --- | --- | --- | --- |
| NA | NA | NA | NA | NA |

  - Safety and performance related information on the in vitro diagnostic medical device:

- Summary of reportable events and field safety corrective action from the date of introduction.

For adverse event (false diagnosis or any other hazard during its use).

| Adverse event (false diagnosis) | Frequency of occurrence during the period (number of reports/total units sold) |
| --- | --- |
| NA | NA |

No such incidence/event is reported for any field safety corrective action from the date of introduction of this in-vitro medical device. The same has been confirmed from the notifying body CDSCO.

For Field Safety Corrective Action (FSCA)

| Date of FSCA | Reason For FSCA | Countries where FSCA was conducted | Description of the action taken |
| --- | --- | --- | --- |
| NA | NA | NA | NA |

- If the in vitro diagnostic medical device contains any of the following, then descriptive information on the following need to be provided.

- Animal or human fluids or derivatives thereof, rendered non-viable.

This is an in vitro diagnostic medical device contains inactivated recombinant antigens and monoclonal antibodies which is not derived from animal origin.

- Cells, tissues and/or derivatives of microbial recombinant origin.

This is an in vitro diagnostic medical device contains inactivated recombinant antigens and monoclonal antibodies which is not derived from animal origin.

Description and specification, including variants and accessories of the in vitro diagnostic medical device.

  - Description

The device master file should include the following device descriptive information:

- It may include.

- What is detected.

The Q-Line® Albumin Assay is designed to quantitatively detect serum albumin, the most abundant plasma protein synthesized by the liver. Albumin plays a critical role in maintaining oncotic pressure, transporting hormones, fatty acids, and drugs, and serving as a marker for nutritional and liver function status. The assay specifically measures the concentration of albumin in biological samples such as serum or plasma, providing valuable clinical information for the diagnosis and monitoring of liver disease, kidney disorders, malnutrition, and other systemic conditions.

- Its function (for example screening, monitoring, diagnostic or aid to diagnosis, staging or aid to staging of disease).

Albumin is a vital protein in human plasma that serves multiple physiological functions. Its primary role is to maintain colloid osmotic (oncotic) pressure, which helps regulate fluid balance between blood vessels and tissues. In addition, albumin acts as a carrier protein, binding and transporting various substances such as hormones, bilirubin, fatty acids, calcium, and therapeutic drugs. It also has a buffering capacity, helping to maintain blood pH, and possesses antioxidant properties by binding free radicals and reactive molecules. Clinically, albumin levels are an important indicator of liver function, nutritional status, and renal health.

- Whether it is automated or not.

This IVD device is an automated device.

- Whether it is qualitative or quantitative.

It is quantitative.

The optimum sample is non haemolyzed fresh serum. Samples should be analysed within one hour at room temperature.

- A general description of the principle of the assay method.

Principle of Assay Method

The Q-Line® Albumin Assay is based on the colorimetric principle using Bromocresol Green (BCG) dye binding. At an acidic pH (approximately 4.2), albumin in the sample binds specifically to the BCG dye to form a green-coloured albumin-BCG complex. This complex exhibits a characteristic absorbance, which is directly proportional to the concentration of albumin in the sample.

The intensity of the color formed is measured spectrophotometrically at 630 nm. The concentration of albumin is determined by comparing the absorbance of the sample to that of a known calibrator or standard.

Chemical Reaction:

Albumin + Bromocresol Green (acidic pH) ⟶ Albumin–BCG Green Complex
(Absorbance measured at 630 nm)

- A description of the specimen collection and transport materials provided with the in vitro diagnostic medical device, or descriptions of specifications recommended for use.

No such specimen collection and transport material are provided with this In-Vitro Medical device.

- For instruments of automated assays; a description of the appropriate assay characteristics or dedicated assays.

It is not used with any automated instrument.

- A description of any software to be used with the in vitro diagnostic medical device.

No software is used with this in vitro diagnostic medical device.

- A description or complete list of the various configurations/variants of the in vitro diagnostic medical device that will be made available.

This In vitro medical device is available in different pack sizes of 100 test per kit.

- A description of the accessories, other in vitro diagnostic medical device and other products that are not in vitro diagnostic medical device, which are intended to be used in combination with the in vitro diagnostic medical device.

Apart from this In vitro medical device (IVD) a R1 Buffer (R1), R2 Buffer (R2), Calibrator, Control Product Insert is provided.

Materials / accessories required but not provided.

2.2 	For a new in vitro diagnostic medical device

Where relevant to demonstrating conformity to the essential principles, and to provide general background information, the device master file may provide a summary of Clinical Performance Evaluation reports.

This is not a new In-vitro medical device. But the documents are in conformity to the essential principles check list and provides general background information regarding the device. As this, product comes under Class-B category, the performance evaluation is not required by notified body.

2.3 	For an in vitro diagnostic medical device already available on the market in India.

- This information may include a summary of the number of adverse event reports related to the safety and performance of this in vitro diagnostic medical device in relation to the number of in vitro diagnostic medical devices placed on the market.

There is no such incidence happen in the market related to this product as evident from CDSCO which is a notifying body in this regard.

- External certificates and documents which give written evidence of conformity with the essential principles may be annexed to the device master file.

Attached: Certificate of registration for ISO 13485:2016.

- Comparative analysis to prove substantial equivalence to the predicate device(s), if claimed in the application.

No such claim is made in this application for substantial equivalence to predicate this device.

Essential Principles Checklist

- The device master file should include an essential principles checklist that identifies:

- The essential principles of safety and performance.

- Whether each essential principle applies to the in vitro diagnostic medical device and if not, why not.

- The method used to demonstrate conformity with each essential principle that applies.

- the reference to the actual technical documentation that offers evidence of conformity with each method used.

- The method used to demonstrate conformity may include one or more of the following.

- Conformity with recognized or other standards.

- Conformity with a commonly accepted industry test method (reference method).

- Conformity with appropriate in-house test methods that have been validated and verified.

- Comparison to an in vitro diagnostic medical device already available on the market.

- The essential principles checklist should include a cross-reference to the location of such evidence both within the full technical documentation held by the manufacturer and within the Device master file (when such documentation is specifically required for inclusion in the Summary Technical Documentation as outlined in this guidance).

Essential Principles General Requirements

| No. | Essential Requirement | Applies (Yes/No/NA) | Applicable Std/ Procedure | Response |
| --- | --- | --- | --- | --- |
| 1 | The Reagent must be designed and manufactured in such a way that when used under the conditions and for the purpose intended, they will not compromise, directly or indirectly. The clinical condition or the safety of patients or the safety and health of users or, where applicable. Other person, or the safety of property. Any risks which may be associated with their use must be acceptable when weighed against the benefits to the patient and be compatible with a high level of protection of health and safety. | Yes | EN ISO 14971:2012 EN 13485:2016 EN 23640:2015 EN 13612:2002 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-2:2011 | Risk Analysis Record, Drawings and Specification Package Insert |
| 2 | The solutions adopted by the manufacturer for the design and construction of the Reagents must conform to safety principles, taking account of the generally acknowledged state of the art. In selecting the most appropriate solutions, the manufacturer must apply the following principles in the following order. - eliminate or reduce risks as far as possible (inherently safe design and construction) -Where appropriate take adequate protection measures including alarms, if necessary, in relation to risks that cannot be eliminated. Inform users of the residual risks due to shortcomings of the protection measures adopted. |  | EN ISO 14971:2012 EN 13485:2016 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-2:2011 | Instructions For Use. Device and procedures for use and preparation are described in the package insert |
| 3 | The Reagents must be designed and manufactured in such a way that they are suitable for the purposes referred to in Article 1(2) (b), as specified by the manufacturer, taking account of the generally acknowledged state of the art. Repeatability and reproducibility of known relevant interference, and limits of detention stated by the manufacturer. The traceability of must be assured through available reference measurement procedures and/or available reference materials of a higher order. |  | EN 23640:2015 EN 13612:2002 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-2:2011 | Performance of device was tested for repeatability, reproducibility, and Limitations. Testing records in Design History File. |
| 4 | The characteristics and performance referred to in sections 1 and 3 must not be adversely affected to such a degree that the health or the safety of the patients or the user and, where applicable, of other persons are compromised during the lifetime of the reagent as indicated by the manufacturer, when the reagent is subjected to the stresses which can occur during normal conditions of use. When no lifetime is stated, the same applies for the lifetime reasonably to be expected of a Reagent of that kind having regard to the intended purpose and the anticipated use of the Reagent. |  | EN 23640:2015 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-2:2011 | Stability data is included in the file. Stability testing records in Design History File. Stability data supports the performance of the Reagent during the lifetime of product. |
| 5 | The Reagents must be designed, manufactured and packed in such a way that their characteristics and performances during their intended use will not be adversely affected during transport and storage (temperature, humidity, etc.) taking into account the instructions and information provided by the manufacturer. |  | EN23640 Stability testing | Production SOPs in File. Stability data is included in File. Reagents are manufactured and packaged in clean condition. The final product is kept in a closed sealed pouch until used. Stability studies at different temperatures ensure stability during the shipment through different temperatures. |

Design and Manufacturing Requirements

| No | Essential Requirement | Applies (Yes/No/NA) | Applicable Std /Procedure | Response |
| --- | --- | --- | --- | --- |
| 1. 1.1 | Physical and Chemical The Reagents must be designed and manufactured in such a way as to achieve the characteristics and Performances referred to in Section A on the "General Requirements". Particular attention must be paid to the possibility of impairment of analytical performance intended to be used with the Reagent, taking account of its intended purpose. |  | EN ISO 14971:2012 EN 13485:2016 EN 23640:2015 EN 13612:2002 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | Potentially interfering substances and potentially cross-reacting endogenous proteins were tested and listed in Instruction for Use. Testing data in File |
| 1.2 | The Reagents must be designed, manufactured, and packed in such a way as to reduce as far as possible the risk posed by product leakage, contaminants and residues to the persons involved in the transport, storage and use of the Reagents, taking account of the intended purpose of the product. | Yes | EN ISO 14971:2012 | All devices are manufactured as per the required formulation and tested and sealed in closed pouch. |
| 2. 2.1 | Infection and Microbial Contamination The reagents and their manufacturing processes must be designed in such a way as to eliminate or reduce as far as possible the risk of infection to the user or other persons. The design must allow easy handling and, where necessary, reduce as far as possible contamination of, and leakage from, the Reagent during use. The manufacturing processes must be appropriate for these purposes. | No | EN ISO 14971:2012 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | No infectious materials used in the production |
| 2.2 | Reagents do not in corporate biological substances, there is no risk of infection. | No | EN ISO 14971:2012 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | No Biological substance is used in the reagent. |
| 2.3 | Reagents labelled either as 'STERILE' or as having a special microbiological state must be designed, manufactured and packed in an appropriate pack, according to procedures suitable for ensuring that they remain in the appropriate microbiological state indicated on the label when placed on the market, under the storage and transport conditions specified by the manufacturer, until the protective packaging is damaged or opened. | No | ISO 11737-2. 2019 (en) | The Device is not a sterile product. |
| 2.4 | Reagent labelled either as 'STERILE' or as having a special microbiological state must have been processed by an appropriate, validated method. | No | ISO 11737-2. 2019 (en) | Device is not, Sterile product. |
| 2.5 | Packaging systems for reagent other than those referred to in section 2.3 must keep the product without deterioration at the level of cleanliness indicated by the manufacturer and, if the Reagents are to be sterilized prior to use, reduce as far as possible the risk of microbial contamination. Steps must be taken to reduce as far as possible microbial contamination during selection and handling of raw materials, manufacture, storage, and distribution where the performance of the Reagent can be adversely affected by such contamination. | Yes | EN ISO 14971:2012 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | All the material in the device is dry and free of moisture, sealed and packed in low relative humidity so that there is no risk of microbial contamination /or it is minimal. Stability data is included in File. |
| 2.6 | Reagents intended to be sterilized must be manufactured in appropriately controlled (e.g., environmental) conditions. | Yes |  | Device is not a sterile product. |
| 2.7 | Packaging systems for non-sterile Reagents must keep the product without deterioration in the level of cleanliness stipulated and, if the Reagents are to be sterilized prior to use, minimize the risk of microbial contamination. the packaging system must be suitable taking account of the method of sterilization indicated by the manufacturer. | Yes | EN ISO 14971:2012 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | The device is packed under low temperature and low relative humidity to minimize the risk of microbial contamination. |
| 3 3.1 | Manufacturing and Environmental If the Reagent is intended for use in combination with other Reagents or equipment, the whole combination, including the connection system must be safe and must not impair the specified performances of the reagents. Any restrictions on the use must be indicated on the label or in the instructions for use. | NA | EN ISO 14971:2012 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | The device is not used in combination with other equipment or Reagent. All necessary buffers and accessories are kept in kit box. |
| 3.2 | Reagents must be designed and manufactured in such a way as to reduce as far as possible the risks linked to their use in conjunction with materials, substances and gases with which they may come into contact during normal conditions of use. |  | EN ISO 14971:2012 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | The device is not used in combination with other equipment or Reagent. |
| 3.3 | Reagents must be designed and manufactured in such a way as to remove or reduce as far as possible. — The risk of injury linked to their physical features (aspects of volume x pressure, dimension and, where appropriate, ergonomic features). — Risks linked to reasonably fore seeable external influences, such as magnetic fields, external electrical effects, electrostatic discharge, pressure, humidity, temperature or variations in pressure or acceleration or accidental penetration of substances into the Reagent. Reagents must be designed and manufactured in such a way as to provide an adequate level of intrinsic immunity to electromagnetic disturbance to enable them to operate as intended. | Yes es NA | EN ISO 14971:2012 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | The device does not cause any physical injury to users. The device is sealed in a single laminated pouch containing silica gel to control moisture. Device is not affected by, Electromagnetic Disturbance. |
| 3.4 | Reagents must be designed and manufactured in such a way as to reduce as far as possible the risks of fire or explosion during normal use and in single fault condition. Particular attention must be paid to Reagents whose intended use includes exposure to or use in association with flammable substances or substances which could cause combustion. | NA | EN ISO 14971:2012 EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | Device is not. used with any flammable substance |
| 3.5 | Reagents must be designed and manufactured in such a way as to facilitate the management of safe waste disposal. | Yes | EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | Device can be disposed of safely as Bio-waste management. |
| 3.6 | The measuring, monitoring, or display scale (including colour change and other visual indicators) must be designed and manufactured in line with ergonomic principles, taking account of the intended purpose of the Reagent. | NA | EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | The device is used manually without an instrument. Operator can use the device easily without any strain of body position. |
| 4 4.1 | Reagents which are instruments or apparatus with measuring function. Reagents which are instruments or apparatus having a primary analytical measuring function must be designed and manufactured in such a way as to provide adequate stability and accuracy of measurement within appropriate accuracy limits, considering the intended purpose of the Reagent and of available and appropriate reference measurement procedures and materials. The accuracy limits must be specified by the manufacturer. | NA | EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-2:2011 EN 13640:2002 EN 13612:2002 | Device does not have analytical measuring function. |
| 4.2 | When values are expressed numerically, they must be given in legal units conforming to the provision of Council Directive 80/181/EEC of 20 December 1979 on the approximation of the laws of the Member States relating to units of measurement. | NA | EN 980:2008 EN ISO 18113-1:2011 EN ISO 18113-:2011 | Device does not express numerical value. |
| 5 5.1 | Protection Against Radiation Reagents shall be designed, manufactured, and packaged in such a way that exposure of users and other persons to the emitted radiation is minimized. | NA |  | Device does not emit radiation. |
| 5.2 | When Reagents are intended to emit potentially hazardous, visible and/or invisible radiation, they must as far as possible be designed and manufactured in such a way as to ensure that the characteristics and the quantity of radiation emitted can be controlled and/or adjusted, Fitted with visual displays and/or audible warnings of such emissions. | NA |  | Device does not emit hazardous visible or invisible radiation. |
| 5.3 | The operating instructions for Reagents emitting radiation must give detailed information as to the nature of the emitted radiation, means of protecting the user, and on ways of avoiding misuse and of eliminating the risks inherent in installation. | NA |  | Device does not emit radiation |
| 6 6.1 | Requirements for medical Reagents connected to / or equipped with an energy source. Reagents incorporating electronic programmable systems, including software, must be designed to ensure the repeatability, reliability, and performance of these systems according to the intended use. | NA |  | Device does not require electronic programmable system. |
| 6.2 | Reagents must be designed and manufactured in such a way as to minimize the risks of creating electromagnetic perturbation which could impair the operation of other Reagents or equipment in the usual environment. | NA |  | Device does not use electricity, therefore, does not emit any electromagnetic wave |
| 6.3 | Reagents must be designed and manufactured in such a way as to avoid, as far as possible, the risk of accidental electric shocks during normal use and in single fault condition, provided the Reagents are installed and maintained correctly. | NA |  | Device does not use electricity. |
| 6.4 | Protection against mechanical and thermal risks |  |  | . |
| 6.4.1 | Reagents must be designed and manufactured in such a way as to protect the user against mechanical risks. Reagents must be sufficiently stable under the foreseen operating conditions. They must be suitable to withstand stresses inherent in the foreseen working environment, and to retain this resistance during the expected life of the Reagents, subject to any inspection and maintenance requirements as indicated by the manufacturer. Where there are risks due to the presence of moving parts, risks due to break-up or detachment, or leakage of substances, then appropriate protection means must be incorporated. Any guards or other means included with the Reagent to provide protection, against moving parts, must be secure and must not interfere with access for the normal operation of the Reagent, or restrict routine maintenance of the Reagent as intended by the manufacturer. | Yes NA NA | EN ISO 14971:2012 EN 23640:2015 EN 980:2008 EN ISO 18113-:2011 EN ISO 18113-:2011 | The size of the device is not enough to cause any physical damage to user. Storage temperature is mentioned on the label of pouches / outer box and in pack insert. There is no risk due to leakage. But the risk is due to tearing of pouch, therefore, to prevent moisture absorption by device silica gel is incorporated. No moving part is included in the reagent. |
| 6.4.2 | Reagents must be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the Reagents, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. | Yes |  | Device is secured in aluminium sealed pouch with silica gel. |
| 6.4.3 | Reagents must be designed and manufactured in such a way as to reduce as far as possible the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. | NA |  | Device does not emit any noise. |
| 6.4.4 | Terminals and connectors to electricity, gas, or hydraulic and pneumatic energy supplies which the user has to handle must be designed and manufactured in such a way as to minimize all possible risks. | NA |  | Device does not require gas or hydraulic and pneumatic energy. |
| 6.4.5 | Accessible parts of the Reagents (excluding the parts of areas intended to supply heat or reach given temperatures) and their surroundings must not attain potentially dangerous temperatures under normal use. |  | EN ISO 14971:2012 EN 23640:2015 EN 980:2008 EN ISO 18113-:2011 | No accessible parts are involved with the device. |
| 7 | Requirement for Reagents for self-testing Reagents for self-testing must be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to users and the influence resulting from variation that can reasonably be anticipated in users' technique and environment. The information and instructions provided by the manufacturer should be easily understood and applied by the user. | NA |  | The device is for professional use only. |
| 7.1 | Reagents for self-testing must be designed and manufactured in such a way as to. — ensure that the Reagent is easy to use by the intended lay user at all stages of the procedure and — Reduce as far as practicable the risk of user error in handling of Reagent and in interpretation of results. | NA |  | The device is for professional use only. |
| 7.2 | Reagents for self-testing must, where reasonably possible, include user control, i.e., a procedure by which the user can verify that, at the time of use the product will perform as intended. | NA |  | The device is for professional use only. |
| 8 8.1 | Information supplied by the manufacturer. Each Reagent must be accompanied by the information needed to use it safely and properly, taking account of the training and knowledge of the potential users, and to identify the manufacturer. This information comprises the data on the label and the instructions for use. As far as practicable and appropriate, the information needed to use the Reagent safely and properly must be set out on the Reagent itself and/or, where appropriate, on the sales packaging. If individual full labelling of each unit is not practicable, the information must be set out on the packing and/or in the instructions for use supplied with one or more Reagents. Instructions for use must accompany or be included in the packaging of one or more Reagents. In duly justified and exceptional cases no such instructions for use are needed for a Reagent if it can be used properly and safely without them. | Yes | EN ISO 14971:2012 EN 980:2008 ISO 15223-1-2012 EN ISO 18113-1:2011 EN ISO 18113-:2011 | Pack Insert (Instruction for use) is provided in each kit box. The content of Pack insert is Instruction for use. Performance data, contact number and address of manufacturer. It is written in simple language |
| 8.2 | Where appropriate, the Information to be supplied should take the form of symbols. Any symbol and identification colour used must conform to the harmonized standards. In areas for which no standards exist, the symbols and colour used must be described in the documentation supplied with the Reagent | Yes | EN ISO 14971:2012 EN 980:2008 ISO 15223-1-2012 EN ISO 18113-1:2011 EN ISO 18113-:2011 | Instruction For Use in Section #4 Technical File. Labelling in Section #8 Technical File. Conform to the harmonized standards. |
| 8.3 | In the case of Reagents containing or a preparation which may be considered as being dangerous, taking account of the nature and quantity of its constituents and the form under which, they are present, relevant danger symbols and labelling Requirements of Directive 67/548/EEC and Directive 88/379/EEC shall apply. Where there is insufficient space to put all the information on the Reagent itself or on its label, the relevant danger symbols shall be put on the label and the other information required by those Directives shall be given in the instructions for use. The provisions of the Directives on the safety data sheet shall apply unless all relevant information as appropriate is already made available by the instructions for use. |  | EN ISO 14971:2012 EN 980:2008 ISO 15223-1-2012 EN ISO 18113-1:2011 EN ISO 18113-:2011 | Device is small and does not contain any dangerous Components. Device is considered as not being dangerous. Device Specification and Device Drawing in File. |
| 8.4 | The label must bear the following particulars which may take the form of symbols as appropriate. - The name or trade name and address of the manufacturer. For Reagents imported into the Community with a view to their distribution in the Community, the label, the outer packaging, or the instructions for use shall contain in addition the name and address of the authorized representative of the manufacturer. - The details strictly necessary for the user to uniquely identify the Reagent and the contents of the packaging. - Where appropriate, the word 'STERILE' or a statement indicating any special microbiological state or state of cleanliness. - The batch code, preceded by the word 'LOT', or the serial number. - If necessary, an indication of the date by which the reagent or part of it should be used, in safety, without degradation of performance, expressed as the year, the month and, where relevant, the Days, in that order. - In case of reagents for performance evaluation, the words 'for performance evaluation only. - Where appropriate, a statement indicating the in vitro use of the reagent. - Any storage and/or handling conditions. - Where applicable, any operating instructions. - Appropriate warnings and/or precautions to take. - If the Reagent is intended for self-testing, that fact must be clearly stated. | Yes Yes Yes Yes Yes Yes Yes Yes Yes No | EN 980:2008 ISO 15223-1-2012 EN ISO 18113-1:2011 EN ISO 18113-:2011 EN ISO 15223-1 Labeling | Instructions For Use. Trade name on pouch, outer box and pack insert. Name/address of company, POCT Services Pvt. Ltd., H-2, Sec-5 Bawana, on pack insert and box. Test Name, catalogue No. and content on the box and/or pack insert. Yes, it is indicated as a NON-Sterile Product. Lot number on box. Expiration date on box Yes, the word performance evaluation is used. IVD is mentioned. Mentioned on label, outer box. Mentioned in pack insert. Mentioned in pack insert. Not for self-testing. The device should be used by trained professionals only. |
| 8.5 | If the intended purpose of the Reagent is not obvious to the user, the manufacturer must clearly state the intended purpose in the instructions for use and, if appropriate, on the label. | Yes | EN ISO 15223-1 labelling EN ISO 18113-2 Instruction for use | Instructions For Use. Intended purpose of the device is stated in the instruction of use. |
| 8.6 | Wherever reasonable and practicable, the Reagents and separate components must be identified, where appropriate in terms of batches, to allow ail appropriate action to detect any potential risk posed by the Reagents and detachable components. |  | EN ISO 14971:2012 EN 980:2008 ISO 15223-1-2012 EN ISO 18113-1:2011 EN ISO 18113-:2011 | Labelling in Section, technical File. Production record of each lot (batch) is kept in Batch Manufacturing Record (BMR). |
| 8.7 | Where appropriate, the instructions for use must contain the following particulars. - The details referred to in section 8.4 except for points d) and e). - Composition of the reagent product by nature and amount or concentration of the active (ingredient(s) of the reagent (s) or kit as well as a statement, where appropriate, that the Reagent contains other ingredients which might influence the measurement. - The storage conditions and shelf life following the first opening of the primary container, together with the storage conditions and stability of working reagents. - The performances referred to in section 3 of part A. - An indication of any special equipment required including information necessary for the identification of that special equipment for proper use. - The type of specimen to be used, any special conditions of collection, pre-treatment and, if necessary, storage conditions and instructions for the preparation of the patient. - A detailed description of the procedure to be followed in using the Reagent. - The measurement procedure to be followed with the Reagent including as appropriate. —The principle of the method. —The specific analytical performance characteristics (e.g., sensitivity, specificity, accuracy, repeatability, reproducibility, limits of detection and measurement range, including information needed for the control of known relevant interferences), limitations of the method and information about the use of available reference measurement procedures and materials by the user. The details of any further procedure or handling needed before the reagent can be used (for example, reconstitution, incubation, dilution, instrument checks, etc.) An indication whether any training is required. - The mathematical approach upon which the calculation of the analytical result is made. - Measures to be taken in the event of changes in the analytical performance of the reagent. - Information appropriate to users on. a. internal quality control including specific validation procedures. b. The traceability of the calibration of the Reagent. - The reference intervals for the quantities being determined, including a description of the appropriate reference population. - If the Reagent must be used in combination with or installed with or connected to other medical Reagents or equipment to operate as required for its intended purpose, sufficient details of its characteristics to identify the correct Reagents or equipment to use to obtain a safe and proper combination. - All the information needed to verify whether the Reagent is properly installed and can operate correctly and safely, plus details of the nature and frequency of the maintenance and calibration needed to ensure that the Reagent operates properly and safely. information about safe waste disposal. - Details of any further treatment or handling needed before the Reagent can be used (for example, sterilization, final assembly, etc.). - The necessary instructions in the event of damage to the protective packaging and details of appropriate methods of re-sterilization or decontamination. - If the Reagent is reusable, information on the appropriate processes to allow reuse. Including cleaning, disinfecting, packaging and re-sterilization or decontamination, and any restriction on the number of reuses. - Precautions to be taken as regards exposure, reasonably foreseeable environmental conditions, magnetic fields, external electrical influences, electrostatic discharge, pressure or variations in pressure, acceleration, thermal ignition sources, etc. - Precautions to be taken against any special, unusual risks related to the use or disposal of the Reagent including special protective measures, where the Reagent includes substances of human or animal origin, attention must be drawn to their potential infectious nature. - Specifications for Reagents for self-testing. The results need to be expressed and presented in a way that is readily understood by a lay person. Information needs to be provided with advice to the user on action to be taken (in case of positive, negative, or indeterminate result) and on the possibility of false positive or false negative result. - Specific particulars may be omitted provided that the other Information supplied by the manufacturer is sufficient to enable the user to use the Reagent and to understand the result (s) produced by the Reagent. - The information provided must include a statement clearly directing that the user should not take any decision of medical relevance without first consulting his or her medical practitioner. - The information must also specify that when the Reagent for self-testing is used for the monitoring of an existing disease, the patient should only adapt the treatment if he has received the appropriate training to do so. - Date of issue or latest revision of the instructions for use. | Yes Yes Yes Yes NA Yes Yes Yes NA NA Yes Yes NA NA NA Yes NA NA NA NA Yes Yes NA | EN 980:2008 ISO 15223-1-2012 EN ISO 18113-1:2011 EN ISO 18113-:2011 EN ISO 15223-1 Labeling | Instruction of Use. Instructions For Use. Lot number and expiration date. Pack inserts and/or box. Storage temperature and expiration date on label/ box. Pack inserts. No special instrument is required. Pack inserts. Pack inserts. Interpretation of results mentioned in pack insert. No calculation needed. Not analytical performance Lot number of product can be used to trace the Albumins. Mentioned in pack insert. The device is used alone. Installation is not required. No other steps need before being used. Non-Sterile Product The device is not reusable. Device is not affected by magnetic fields, external electrical influences, and electrostatic discharge. Etc. No human substance. The device is for professionals. use only. Information supplied by the manufacturer is sufficient for user to use the device and to understand the results produced by the reagent. The device is IVD device, and it is clearly mentioned not to use without consulting medical practitioner. The existing device is not for self-testing. Part number with revision number in the pack insert. |

Risk Analysis and Control Summary.

The device master file should contain a summary of the risks identified during the risk analysis process and a description of how these risks have been controlled to an acceptable level. Preferably, this risk analysis should be based on recognised standards and be part of the manufacturer’s risk management plan.

The summary should address possible hazards for the in vitro diagnostic medical device such as the risk from false positive or false negative results, indirect risks which may result from in vitro diagnostic medical device associated hazards, such as instability, which could lead to erroneous results, or from user-related hazards, such as reagents containing infectious agents. The results of the risk analysis should provide.

a conclusion with evidence that remaining risks are acceptable when compared to the benefits.

- The Risk Assessment Matrix

A risk assessment matrix is a tool that was developed to analyse risk. Yes, we can use data to analyse risks. By doing so, any organization can detect and prioritize different risks. They do this by estimating the probability of occurrence.

Figure 6.1: Risk Assessment Matrix

| Risk Matrix | Minor | Moderate | Major | Critical |
| --- | --- | --- | --- | --- |
| 76-100% | 0 | 0 | 0 | 2 |
| 51-75% | 2 | 0 | 5 | 0 |
| 26-50% | 0 | 3 | 0 | 0 |
| 0-25% | 1 | 12 | 0 | 0 |

A risk matrix is sometimes also called the Probability Matrix, or Impact Matrix. This is an effective tool that can help in risk evaluation by focusing on the probability of potential risks. A risk assessment matrix can help to calculate Product risk quickly. It does this by identifying the things that could go wrong and weighting the potential damage. This makes it easy to prioritize problems. Action will be needed in order to keep a Product on course, and safe as well.

Product managers should think about potential risks in order to avoid risk events from happening. While managing uncertainty sounds challenging, there are more and more calculating risk tools available to Days that can help and require little effort on your part. Simply create your own risk assessment matrix and use it as many times as you need.

Benefits of When Making Risk Matrix.

- You will be able to prioritize the risks with the level of severity.

- You get a simple process for the management of risk.

- It helps you in finding the potential risk with minimal effort.

- Information is recorded and audited.

- It demonstrates the organization’s ability to managing risk.

- It helps in neutralizing any possible consequences.

- How To Make a Risk Assessment Matrix.

If you want to do your own risk assessment matrix, you can start by defining the scope of work. Depending on what you are trying to improve, you need to identify different areas of risk. Choose your objective and make sure it is clear as possible.

Figure 6.2: Identify Hazards

| SEVERITY | Risk Assessment Matrix |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | Consequences |  |  |  | Increasing Likelihood |  |  |  |  |
|  | People | Assets | Community | Environment | A | B | C | D | E |
|  |  |  |  |  | Never heard of in the industry | Heard of in the industry | Has happed in our organization or more than once per year in the industry | Has happed at the location or more than once per year in our organization | Has happed more than once per year at location |
| 0 | No injury or health effect | No damage | No effect | No effect | Continuous Improvements |  |  |  |  |
| 1 | Slight injury or health effect | Slight damage | Slight effect | Slight effect |  |  |  |  |  |
| 2 | Minor injury or health impact | Minor damage | Minor effect | Minor effect |  |  |  | Control to ALARP |  |
| 3 | Major injury or health effect | Moderate damage | Moderate effect | Moderate effect |  |  |  |  |  |
| 4 | Permanent total disability or upto three fatalities | Major damage | Major effect | Major effect |  |  |  |  |  |
| 5 | More than three fatalities | Massive damage | Massive effect | Massive effect |  |  | Tolerability to be endorsed by management |  |  |

To start, you want to go for as many risks as you can. The idea behind this is to get different views. A brainstorming session could be of help. The list that you get is going to be the foundation of the risk assessment matrix.

Connected with your scope, the list needs to belong and detailed. It can include anything from theft, to burns, and even pollution. It is important that you think at all potential risks for any new Product you are working on.

You can also think about what happens when you identify them. But not to worry, we will discuss that soon enough.

Figure 6.3: Risk Analysis

| Hazard analysis and risk assessment matrix with proposed actions |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| HAZARD SEVERITY |  |  |  |  |  |
| Hazard likelihood | Critical Illness or Injury | Severe Illness or Injury | Moderate Illness or Injury | Minor Impact | Negligible Impact |
| Very Likely | Requires Controls | Requires Controls | Requires Controls | Manageable | Manageable |
| Likely | Requires Controls | Requires Controls | Manageable | Manageable and Tolerable | Tolerable |
| Possible | Requires Controls | Manageable | Manageable and Tolerable | Tolerable | Acceptable |
| Unlikely | Manageable | Tolerable | Tolerable | Acceptable | Acceptable |
| Highly Unlikely | Tolerable | Acceptable | Acceptable | Acceptable | Acceptable |

The risk analysis is not something to take lightly. There are certain steps that you need to follow to do effective management of risks. When an organization has pitched all the right risks, the next step is going to carefully evaluate them. A risk assessment matrix focuses a lot of chances and consequences as the focus. But depending on the organization, we are talking about you can encounter terms like “vulnerability” or “speed of onset”.

Figure 6.4: Determining Risk Impact

| Likelihood of residual risk | Almost Certain 5 | 5 Supplementary Issue | 10 Issue | 15 Unacceptable | 20 Unacceptable | 25 Unacceptable |
| --- | --- | --- | --- | --- | --- | --- |
|  | Probable 4 | 3 Acceptable | 8 Supplementary Issue | 12 Issue | 16 Unacceptable | 20 Unacceptable |
|  | Possible 3 | 3 Acceptable | 6 Supplementary Issue | 9 Issue | 12 Issue | 15 Unacceptable |
|  | Unlikely 2 | 2 Acceptable | 3 Acceptable | Supplementary Issue | 8 Supplementary Issue | 10 Issue |
|  | Rare 1 | 1 Acceptable | 2 Acceptable | 3 Acceptable | 4 Acceptable | 5 Issue |
|  |  | Insignificant 1 | Minor 2 | Moderate 3 | Major 4 | Catastrophic 5 |
|  | Consequence |  |  |  |  |  |

Any risk assessment matrix means that you will need to check probabilities and consequences of risk events that might happen. The results of such assessments are used to make a top of risks to find the most important ones, as well as less critical ones.

In a risk chart, you can see exactly how both high-risk and low-risk factors are shown. The impact of a successful attack can be split into two types. “Technical impact” and the “business impact”.

Step 4. Prioritize the risk.

When you will see a risk assessment matrix, you will be able to compare different levels of risk. It can include any internal rules or policies.

One thing that should be Noted is that the risk assessment process can be an ongoing evolution. A matrix needs to change at the same time with changes that appear in your company. If it is done one timer per year, emerging risks could go unnoticed or even undetected.

How to use the risk assessment matrix?

Figure 6.5: Likelihood. the probability of a risk

| Probability |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  |  | 1 = High (80% ≤ x 100%) | 2 = Medium High (60% ≤ x 80%) | 3 = Medium Low (30% ≤ x 60%) | 4= Low (0% ≤ x 30%) |
| Impact | A = High (Rating 100) | (Exposure -Very High) (Score 100) | (Exposure -Very High) (Score 80) | (Exposure - High) (Score 60) | (Exposure - Moderate) (Score 30) |
|  | B = Medium (Rating 50) | (Exposure - High) (Score 50) | (Exposure - Moderate) (Score 40) | (Exposure- Moderate) (Score 30) | (Exposure - Low) (Score 15) |
|  | C= Low (Rating 10) | (Exposure - Low) (Score 10) | (Exposure - Low) (Score 8) | (Exposure - Low) (Score 6) | (Exposure - Low) (Score 3) |

When the risk assessment process is complete, you can start to take data into the matrix. Any risk assessment matrix uses two axes, one that measures the likelihood, and the other one measures the consequence result.

Depending on the likelihood of the occurrence of the risk, the risk can be classified under these categories.

  - A risk that is almost guaranteed to show up during the execution of the Product. Any risk that is more than 85% likely to cause problems is going to fall under this category.

  - Risks that have a 60%-80% chance to occur can be grouped as likely.

  - Risks that have a 50/50 probability of occurrence are named occasional.

  - Seldom are the risks that have a low probability of occurrence.

  - Unlikely are the risks that have almost no probability of occurring.

Consequences. the severity of the impact or the extent of damage caused by the risk.

Figure 6.6: Likelihood. the probability of a risk

| Likelihood (Probability) How Likely is the event to occur at some time (Linear Scale time specific matrix) | Consequences |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
|  | What is the Severity of injuries / potential damages/financial impacts (If the risk event occurs) (Logarithmic Scale, property industry specific matrix) |  |  |  |  |
|  | Insignificant | Minor | Moderate | Major | Catastrophic |
|  | No Injuries First Aid No Envir Damage << ₹ 1,000 Damage | Some First Aid required Low Envir Damage << ₹ 10,000 Damage | External Medical Medium Envir Damage << ₹ 100,000 Damage | External injuries High Envir Damage << ₹ 1,000,000 Damage | Death or Major injuries Toxic Medium Envir Damage << ₹ 1,0000,000 Damage |
| Almost Certain | Moderate Risk | High Risk | High Risk | Critical Risk | Critical Risk |
| Expected in Normal Circumstances (100 %) |  |  |  |  |  |
| Likely | Moderate Risk | Moderate Risk | High Risk | High Risk | Critical Risk |
| Probably Occur in most Circumstances (10 %) |  |  |  |  |  |
| Possible | Low Risk | Moderate Risk | High Risk | High Risk | Critical Risk |
| Might Occur at Some Time (1 %) |  |  |  |  |  |
| Unlikely | Low Risk | Moderate Risk | Moderate Risk | High Risk | High Risk |
| Could Occur at Some Future Time (0.1 %) |  |  |  |  |  |
| Rare | Low Risk | Low Risk | Moderate Risk | Moderate Risk | High risk |
| Only in Exceptional Circumstances (0.1 %) |  |  |  |  |  |

The consequences of risk can be ranked into five categories. These are based on how severe the damage can get.

- Risks that can cause the negligible amount of damage are called insignificant.

- Risks that have a small potential for negative effects are called minor.

- Risks that do not impose a great threat but are yet sizable damage can be classified as moderate.

- Risks that have substantial negative effects and are going to impact in a serious way the success of a Product is called critically.

Risks that come from human error or the environment. Other causes can be procedural deficiencies or major system loss. This will require the closing on the operation and are called catastrophic.

Knowing what elements, a risk assessment matrix has is important. This is going to help you and your organization to manage risk effectively and reduce workplace incidents. The risk assessment matrix is a document that must be updated and maintained with curiosity. Risks are evolving and the matrix should do the same. There are certain events that are going to trigger the need for a refresh. One could be like establishing an enterprise risk management program.

Ending thoughts on risk assessment matrix

In conclusion, a risk assessment matrix is only going to do good things for you and your organization. Whether this is the first time you are doing it, or you are experienced, you can check the information above and already have some ideas regarding what you need to do. Do not forget that without it you can potentially create a lot of havoc in your company, causing a loss in productivity and time. You will want to study it carefully and see if you can do it independently.

Risk Control.

Risk reduction. Risk control Option analysis.

The risk control measures shall be used one or more followings in the priority order.

Inherent safety by design.

Protective measures in the medical device itself or in the manufacturing process.

Information for safety.

Implementation of risk control measures.

The risk control measures adopted shall be implemented and verified.

Residual risk evaluation.

After the risk control measures are applied, the residual risk evaluation shall be performed in accordance with the same criteria.

Risk/benefit analysis.

Decide if the overall residual risks posed by the medical Reagents are acceptable.

Other generated hazards.

Risk control measures taken in step 5 may have inherent hazard. If any new hazards are introduced by any risk control measures, the associated risk(s) shall be assessed.

Completeness of risk evaluation (Step 10)

The risks from all identified hazards have been evaluated and decided to be acceptable.

| Hazards | Identified Number | Number of Risk control measures | Evaluation |  |  | Acceptable Unacceptable |
| --- | --- | --- | --- | --- | --- | --- |
|  |  |  | BAR | ALAP | IR |  |
| Manufacturing Process | 16 | 16 | 16 | - | - | A |
| Material | 30 | 30 | 30 | - | - | A |
| Biological Hazard | 7 | 7 | 7 | - | - | A |

Table 6.1: Risk Analysis Product. Q-Line® Albumin

| Product Phase | Risk Management Deliverables |
| --- | --- |
| Planning | - Risk Management Plan |
| Design and Development | - System Risk Analysis (hazard identification) - System Risk Evaluation - Risk Assessment (product & process) |
| Design Verification | - Risk Control - Residual Risk Acceptance |
| Design Validation | - Risk Management Report |
| Market Release | - Production & Post-Production Risk Management - Revised Risk Management Report |

Risk management deliverables are reviewed and approved during design reviews for each Product phase.  ISO 14971.2007 shall be used for instructions and as guidelines during risk management documentation. Refer to Risk Management Procedure for the company process.

Risk is defined as the combination of occurrence of harm and the severity of that harm. To estimate risks of hazardous situations relating to IVD medical devices, severity of harm and probability of occurrence of harm are estimated according to the tables below.

Table 6.2: Severity Table

| Severity Rating | Description |
| --- | --- |
| Critical | Loss of limb. life-threatening injury |
| Major | Severe, long-term injury. potential disability |
| Serious | Short-term injury or impairment requiring additional medical intervention to correct (e.g. reoperation) |
| Minor | Slight customer inconvenience. little to no effect on product performance, non-vital fault |
| Negligible | No or negligible risk to patient |

Table 6.3: Probability Tables

| Probability Rating | Description |
| --- | --- |
| Frequent | 1 in 100 |
| Probable | 1 in 1,000 |
| Occasional | 1 in 10,000 |
| Remote | 1 in 100,000 |
| Improbable | 1 in 1,000,000 |

Risk level is determined based on probability and severity estimates.

Figure 6.7:Risk Estimation

| Frequent | LOW | MEDIUM | HIGH | HIGH | HIGH |
| --- | --- | --- | --- | --- | --- |
| Probable | LOW | MEDIUM | MEDIUM | HIGH | HIGH |
| Occasional | LOW | LOW | MEDIUM | MEDIUM | HIGH |
| Remote | LOW | LOW | LOW | MEDIUM | HIGH |
| Improbable | LOW | LOW | LOW | LOW | MEDIUM |
|  | Negligible | Minor | Serious | Major | Critical |

Risks for all products are identified as LOW, MEDIUM, or HIGH.

Risks in the Negligible, minor shall be acceptable.

Risks in the Serious, major and critical are unacceptable and require risk reduction.

Risk Analysis Associated with This Device:

| Part Name | Part Function | Failure Mode | Cause of Failure | Failure effect | Current Condition |  |  |  | Recommended action | Action Result |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | S | O | D | RPN |  | S | O | D | RPN | Risk level* |
| Tris Buffer (pH 8.0) | Stabilizer for antibody and conjugate. Blocking agent | False positive and/or | Poor quality of raw material | Indicator dye reagent on dye pad was not stable enough. | 5 | 2 | 3 | 30 | Inspect incoming material for lot variation or inspecting coming receiving processor audit supplier certifications | 1 | 1 | 1 | 1 | BA |
|  |  | decreased assay sensitivity |  |  |  |  |  |  |  |  |  |  |  |  |
|  |  |  | Poor quality of BSA solution |  | 5 | 3 | 4 | 60 | Inspect in-process solution preparation or solution handling or solution storage condition. | 1 | 1 | 2 | 2 | BA |
|  |  | Variation of signal intensity | Poor quality of raw material | Quality of dried indicator dye pad was inconsistent. | 4 | 2 | 3 | 24 | Inspect incoming material for lot variation or audit supplier certifications | 1 | 1 | 1 | 1 | BA |
|  |  |  | Poor quality of BSA solution |  | 4 | 3 | 3 | 36 | Inspect in-process sol. | 1 | 1 | 2 | 2 | BA |
|  |  |  |  |  |  |  |  |  | Prep. or sol. handling or solution storage condition. |  |  |  |  |  |
|  |  |  | Poor quality of raw material | Indicator dye pad was pretreated unevenly. | 3 | 2 | 3 | 18 | Inspect incoming material for lot variation or inspect incoming receiving process or audit supplier certifications. | 1 | 1 | 1 | 1 | BA |
|  |  |  |  | Sample application pad was pretreated unevenly. | 3 | 2 | 3 | 18 |  | 1 | 1 | 1 | 1 | BA |
|  |  |  | Poor quality of BSA Solution | Indicator dye pad was Pretreated unevenly. | 3 | 3 | 4 | 36 | Inspect in-process solution preparation or solution handling or solution storage condition. | 1 | 1 | 1 | 1 | BA |

| Part Name | Part Function | Failure Mode | Cause of Failure | Failure effect | Current Condition |  |  |  | Recommended action | Action Result |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  |  |  |  |  | S | 0 | D | RPN |  | S | O | D | RPN | Risk level* |
|  |  |  |  | Sample application pad was pretreated unevenly. | 3 | 3 | 4 | 36 |  | 1 | 1 | 1 | 1 | BA |
| Bromocresol green | Cross linking agent | False negative and/or decreased assay sensitivity | Coupling procedure is not properly performed. | Ag-Ab concentrate signal, not strong enough as designed | 3 | 3 | 3 | 27 | Training personnel for proper procedure | 1 | 1 | 1 | 1 | BA |
|  |  | False positive and/or assay sensitivity increased | Coupling procedure is not properly performed. | Ag-Ab concentrate signal higher than it should be. | 5 | 3 | 3 | 45 | Training personnel for proper procedure | 1 | 1 | 1 | 1 | BA |
| Bovine Albumin | Buffer solution pre-treatment | False positive and/or decreased assay sensitivity | Quality of raw material might not be met specification for the production | Reagents on dye pad and/or membrane were not stable enough. | 5 | 2 | 3 | 30 | Inspection incoming material with CoA | 1 | 1 | 2 | 2 | BA |
|  | and indicator pre- treatment. |  | Buffer solution might be made improperly, preserve poor buffering capacity, or deteriorated. during storage |  | 5 | 3 | 4 | 60 | Training operator for preparation and storage. Check storage condition | 2 | 1 | 2 | 4 | BA |
|  |  | Variation of signal intensity | Buffer solution might be made improperly, preserve poor buffering capacity, or deteriorated during storage | Quality of dried Indicator dye pad was uneven. | 4 | 3 | 4 | 48 | Training operator for preparation and storage. Check storage condition. | 1 | 1 | 2 | 2 | BA |
|  |  | Signal streaked | Buffer solution might be made improperly. | Indicator dye pad was pretreated unevenly. | 4 | 3 | 4 | 48 |  | 1 | 1 | 1 | 1 | BA |
|  |  |  | preserve poor buffering capacity, or deteriorated during storage | Sample pad was pretreated unevenly. | 4 | 3 | 4 | 48 | Training operator for preparation and Check storage condition | 1 | 1 | 1 | 1 | BA |
| XOD | Indicator dye dilution buffer. | Signal streaked | Quality of raw material might not be met specification for the production | Quality of dried Indicator dye pad was uneven. | 4 | 2 | 3 | 24 | Inspection incoming material with CoA | 1 | 1 | 1 | 1 | BA |
|  |  |  | Buffer solution might Be made improperly, preserve poor buffering capacity, or deteriorated during storage |  | 4 | 3 | 4 | 48 | Training operator for preparation and storage. Check storage condition | 1 | 1 | 1 | 1 | BA |
| Brij 35 | Blocking reagent for making sample pre-treatment solution and indicator dye pre-treatment solution. | False positive and/or decreased assay sensitivity | Poor quality of raw material | Reagents on dye pad and/or membrane were not stable enough. | 5 | 2 | 3 | 30 | Inspection incoming material with CoA. | 2 | 1 | 2 | 4 | BA |
|  |  | variation of signal intensity | Poor quality of raw material | Quality of dried indicator dye pad is uneven. | 4 | 2 | 4 | 32 | Inspection incoming material with CoA. | 1 | 1 | 2 | 2 | BA |
|  |  | Signal streaked | Poor quality of raw material | Indicator dye pad was pretreated unevenly. | 4 | 2 | 3 | 24 | Inspection incoming Material with CoA. | 1 | 1 | 1 | 1 | BA |
|  | Keeping device with desiccant. | False positive and/or invalid result | Defected pouch | Membrane strip was deteriorated | 5 | 2 | 3 | 30 | Incoming inspection Standard and sampling plan needs to be evaluated | 1 | 1 | 1 | 1 | BA |
|  |  |  | Pouch Sealing was Not completed |  | 5 | 3 | 4 | 60 | Increase standard of Final QC process and Train operator for sealing pouch | 1 | 2 | 1 | 2 | BA |
|  |  |  | MVTR is higher than desiccant capacity |  | 5 | 2 | 4 | 40 | Redesign pouch material and desiccant capacity. | 1 | 1 | 1 | 1 | BA |
| Stabilizers | Antibody for control line. | Invalid test result (no signal in control band) | Quality of raw material might not be met specification | Reagent for control band was inactive or not sprayed. | 8 | 3 | 3 | 72 | Inspect incoming inspection with CoA according to incoming SOP | 2 | 1 | 1 | 2 | BA |
|  |  |  | Spraying procedure was not handled properly |  | 8 | 3 | 3 | 72 | Check spraying sol & spraying machine. Train operator for spraying procedure. | 2 | 1 | 1 | 2 | BA |
|  |  |  | Antibody was deteriorated during storage. |  | 8 | 3 | 3 | 72 | Check antibody storage condition and check antibody activity before using. | 2 | 1 | 1 | 2 | BA |
| Preservatives | Capturing Antibody in sample | False positive and/or too strong signal intensity | Contamination with other analyte solution during membrane spraying | Non-specific binding to another antibody | 7 | 4 | 5 | 140 | Training operator for membrane spraying procedure. | 2 | 1 | 3 | 6 | BA |
|  |  |  | Quality of incoming material might not meet specification |  | 7 | 2 | 3 | 42 | Incoming inspection with CoA or auditing supplier | 1 | 1 | 2 | 2 | BA |
|  |  |  | Preparation and handling might be improper |  | 7 | 3 | 4 | 84 | Training operator for preparation, handling. | 2 | 1 | 3 | 6 | BA |
|  |  | False negative and/or decreased assay sensitivity | Quality of incoming material might not meet specification | Reagents on dye pad and/or membrane was not stable or specific. | 5 | 2 | 3 | 30 | Incoming inspection With CoA | 1 | 1 | 2 | 2 | BA |
|  |  |  | Preparation and handling might be improper |  | 5 | 3 | 4 | 60 | Training operator for preparation, handling. | 2 | I | 2 | 4 | BA |
| Deionized Water | Holding coated NC Membrane | variation of signal intensity | Drying procedure was not handled properly. | Sheet was dried unevenly. | 4 | 3 | 4 | 48 | Check Freeze drier and train operator to use freeze drier properly. | 2 | 1 | 2 | 4 | BA |
|  |  | variation of signal intensity | Procedure for spraying test band and control band was not handled properly. | Reagents for test band and control band were unevenly coated on the membrane. | 4 | 3 | 4 | 48 | Training operator for spraying procedure and Check performance of spray machine. | 2 | 1 | 3 | 6 | BA |
|  |  |  | Quality of raw material might not be met specification |  | 4 | 2 | 3 | 24 | Incoming inspection standard needs to be evaluated. | 2 | 1 | 2 | 4 | BA |
|  |  | Streaked signal | Procedure for membrane strips was not handled properly. | Membrane strips were not cut evenly. | 4 | 3 | 3 | 36 | Check knife of strip cutter and train operator to use strip cutter properly. | 1 | 1 | 1 | 1 | BA |
|  |  | Invalid test result (Control band did not appear during assay) | Coating machine was not operated properly or was not connected for control band | Reagent for control band was not coated on the membrane. | 8 | 3 | 3 | 72 | Check coating machine and train operator to use coating machine properly. | 2 | 1 | 1 | 2 | BA |
|  | Device in low humidity condition | False positive or invalided result | Desiccant capacity lower than specified | Membrane strip was deteriorated. | 5 | 2 | 3 | 30 | Increase sampling plan and inspection standard for incoming material. | 1 | 1 | 1 | 1 | BA |
| Sodium Azide | Sample movement and background clearance | Assay time longer than specified | Transferring volume per tube is lower than specified | Sample migration rate was slower than designed due to insufficient sample volume. | 4 | 2 | 3 | 24 | Sampling plan and standard for incoming inspection need to be evaluated | 1 | 1 | 2 | 2 | BA |
|  |  |  | Sample transfer was not handled as indicated |  | 4 | 3 | 3 | 26 | Check packages insert to find test procedure description. | 2 | 1 | 1 | 2 | BA |

Design and Manufacturing Information

| No | Essential Requirement | Applies (Yes/No/NA) | Applicable Std /Procedure | Response |
| --- | --- | --- | --- | --- |
| 1. 1.1 | Physical and Chemical The Reagents must be designed and manufactured in such a way as to achieve the characteristics and Performances referred to in Section A on the "General Requirements". Particular attention must be paid to the possibility of impairment of analytical performance intended to be used with the Reagent, taking account of its intended purpose. |  |  | Potentially interfering substances and potentially cross-reacting endogenous proteins were tested and listed in Instruction for Use. Testing data in File |
| 1.2 | The Reagents must be designed, manufactured, and packed in such a way as to reduce as far as possible the risk posed by product leakage, contaminants and residues to the persons involved in the transport, storage and use of the Reagents, taking account of the intended purpose of the product. | Yes |  | All devices are manufactured as per the required formulation and tested and sealed in closed pouch. |
| 2. 2.1 | Infection and Microbial Contamination The reagents and their manufacturing processes must be designed in such a way as to eliminate or reduce as far as possible the risk of infection to the user or other persons. The design must allow easy handling and, where necessary, reduce as far as possible contamination of, and leakage from, the Reagent during use. The manufacturing processes must be appropriate for these purposes. | No |  | No infectious materials used in the production |
| 2.2 | Reagents do not in corporate biological substances, there is no risk of infection. | No |  | No Biological substances are used in the reagent. |
| 2.3 | Reagents labelled either as 'STERILE' or as having a special microbiological state must be designed, manufactured, and packed in an appropriate pack, according to procedures suitable for ensuring that they remain in the appropriate microbiological state indicated on the label when placed on the market, under the storage and transport conditions specified by the manufacturer, until the protective packaging is damaged or opened. | No | ISO 11737-2. 2019 (en) | The Device is not a sterile product. |
| 2.4 | Reagent labelled either as 'STERILE' or as having a special microbiological state must have been processed by an appropriate, validated method. | No | ISO 11737-2. 2019 (en) | Device is not sterile product. |
| 2.5 | Packaging systems for reagent other than those referred to in section 2.3 must keep the product without deterioration at the level of cleanliness indicated by the manufacturer and, if the Reagents are to be sterilized prior to use, reduce as far as possible the risk of microbial contamination. Steps must be taken to reduce as far as possible microbial contamination during selection and handling of raw materials, manufacture, storage, and distribution where the performance of the Reagent can be adversely affected by such contamination. | Yes | EN 23640 Stability testing | All the material in the device is dry and free of moisture, sealed and packed in low relative humidity so that there is no risk of microbial contamination /or it is minimal. Stability data is included in File. |
| 2.6 | Reagents intended to be sterilized must be manufactured in appropriately controlled (e.g., environmental) conditions. | Yes |  | Device is not a sterile product. |
| 2.7 | Packaging systems for non-sterile Reagents must keep the product without deterioration in the level of cleanliness stipulated and, if the Reagents are to be sterilized prior to use, minimize the risk of microbial contamination. the packaging system must be suitable taking account of the method of sterilization indicated by the manufacturer. | Yes |  | The device is packed under low temperature and low relative humidity to minimize the risk of microbial contamination. |
| 3 3.1 | Manufacturing and Environmental If the Reagent is intended for use in combination with other Reagents or equipment, the whole combination, including the connection system, must be safe and must not impair the specified performances of the reagents. Any restrictions on its use must be indicated on the label or in the instructions for use. | NA |  | The device is not used in combination with other equipment or Reagent. All necessary buffers and accessories are kept in the kit box. |
| 3.2 | Reagents must be designed and manufactured in such a way as to reduce as far as possible the risks linked to their use in conjunction with materials, substances, and gases with which they may come into contact during normal conditions of use. |  |  | The device is not used in combination with other equipment or Reagent. |
| 3.3 | Reagents must be designed and manufactured in such a way as to remove or reduce as far as possible. — The risk of injury linked to their physical features (in particular aspects of volume x pressure, dimension and, where appropriate, ergonomic features). — Risks linked to reasonably fore seeable external influences, such as magnetic fields, external electrical effects, electrostatic discharge, pressure, humidity, temperature or variations in pressure or acceleration or accidental penetration of substances into the Reagent. Reagents must be designed and manufactured in such a way as to provide an adequate level of intrinsic immunity to electromagnetic disturbance to enable them to operate as intended. | Yes Yes NA | EN ISO 14971 Risk Management | The device does not cause any physical injury to users. The device is sealed in a single laminated pouch containing silica gel to control moisture. Device is not affected by Electromagnetic Disturbance. |
| 3.4 | Reagents must be designed and manufactured in such a way as to reduce as far as possible the risks of fire or explosion during normal use and in single fault condition. Particular attention must be paid to Reagents whose intended use includes exposure to or use in association with flammable substances or substances which could cause combustion. | NA |  | Device is not used with any flammable substance. |
| 3.5 | Reagents must be designed and manufactured in such a way as to facilitate the management of safe waste disposal. | Yes |  | Device can be disposed of safely as Bio-waste management. |
| 3.6 | The measuring, monitoring, or display scale (including color change and other visual indicators) must be designed and manufactured in line with ergonomic principles, taking account of the intended purpose of the Reagent. | NA |  | The device is used manually without an instrument. Operator can use the device easily without any strain of body position. |
| 4 4.1 | Reagents which are instruments or apparatus with measuring function. Reagents which are instruments or apparatus having a primary analytical measuring function must be designed and manufactured in such a way as to provide adequate stability and accuracy of measurement within appropriate accuracy limits, considering the intended purpose of the Reagent and of available and appropriate reference measurement procedures and materials. The accuracy limits must be specified by the manufacturer. | NA |  | Device does not have analytical measuring function. |
| 4.2 | When values are expressed numerically, they must be given in legal units conforming to the provisions of Council Directive 80/181/EEC of 20 December 1979 on the approximation of the laws of the Member States relating to units of measurement. | NA |  | Device does not express numerical value. |
| 5 5.1 | Protection Against Radiation Reagents shall be designed, manufactured, and packaged in such a way that exposure of users and other persons to the emitted radiation is minimized. | NA |  | Device does not emit radiation. |
| 5.2 | When Reagents are intended to emit potentially hazardous, visible and/or invisible radiation, they must as far as possible be. - designed and manufactured in such a way as to ensure that the characteristics and the quantity of radiation emitted can be controlled and/or adjusted, Fitted with visual displays and/or audible warnings of such emissions. | NA |  | Device does not emit hazardous visible or invisible radiation. |
| 5.3 | The operating instructions for Reagents emitting radiation must give detailed information as to the nature of the emitted radiation, means of protecting the user, and on ways of avoiding misuse and of eliminating the risks inherent in installation. | NA |  | Device does not emit radiation |
| 6 6.1 | Requirements for medical Reagents connected to / or equipped with an energy source. Reagents incorporating electronic programmable systems, including software, must be designed to ensure the repeatability, reliability, and performance of these systems according to the intended use. | NA |  | Device does not require electronic programmable system. |
| 6.2 | Reagents must be designed and manufactured in such a way as to minimize the risks of creating electromagnetic perturbation which could impair the operation of other Reagents or equipment in the usual environment. | NA |  | Device does not use electricity, therefore, does not emit any electromagnetic wave |
| 6.3 | Reagents must be designed and manufactured in such a way as to avoid, as far as possible, the risk of accidental electric shocks during normal use and in single fault condition, provided the Reagents are installed and maintained correctly. | NA |  | Device does not use electricity. |
| 6.4 | Protection against mechanical and thermal risks |  |  | . |
| 6.4.1 | Reagents must be designed and manufactured in such a way as to protect the user against mechanical risks. Reagents must be sufficiently stable under the foreseen operating conditions. They must be suitable to withstand stresses inherent in the foreseen working environment, and to retain this resistance during the expected life of the Reagents, subject to any inspection and maintenance requirements as indicated by the manufacturer. Where there are risks due to the presence of moving parts, risks due to break-up or detachment, or leakage of substances, then appropriate protection means must be incorporated. Any guards or other means included with the Reagent to provide protection, against moving parts, must be secure and must not interfere with access for the normal operation of the Reagent, or restrict routine maintenance of the Reagent as intended by the manufacturer. | Yes NA NA | EN ISO 14971 Risk Management EN ISO 15223-1 Labelling | The size of the device is not enough to cause any physical damage to the user. Storage temperature is mentioned on the label of pouches / outer box and in pack insert. There is no risk due to leakage. But risk is due to tearing of pouch, therefore, to prevent moisture absorption by device silica gel is incorporated. No moving part is included in the reagent. |
| 6.4.2 | Reagents must be designed and manufactured in such a way as to reduce to the lowest possible level the risks arising from vibration generated by the Reagents, taking account of technical progress and of the means available for limiting vibrations, particularly at source, unless the vibrations are part of the specified performance. | Yes |  | Device is secured in aluminum sealed pouch with silica gel. |
| 6.4.3 | Reagents must be designed and manufactured in such a way as to reduce as far as possible the risks arising from the noise emitted, taking account of technical progress and of the means available to reduce noise, particularly at source, unless the noise emitted is part of the specified performance. | NA |  | Device does not emit any noise. |
| 6.4.4 | Terminals and connectors to electricity, gas or hydraulic and pneumatic energy supplies which the user has to handle must be designed and manufactured in such a way as to minimize all possible risks. | NA |  | Device does not require gas or hydraulic and pneumatic energy. |
| 6.4.5 | Accessible parts of the Reagents (excluding the parts of areas intended to supply heat or reach given temperatures) and their surroundings must not attain potentially dangerous temperatures under normal use. |  |  | No accessible parts are involved with the device. |
| 7 | Requirement for Reagents for self-testing Reagents for self-testing must be designed and manufactured in such a way that they perform appropriately for their intended purpose taking into account the skills and the means available to users and the influence resulting from variation that can reasonably be anticipated in users' technique and environment. The information and instructions provided by the manufacturer should be easily understood and applied by the user. | NA |  | The device is for professional use only. |
| 7.1 | Reagents for self-testing must be designed and manufactured in such a way as to. — ensure that the Reagent is easy to use by the intended lay user at all stages of the procedure and — Reduce as far as practicable the risk of user error in handling of Reagent and in interpretation of results. | NA |  | The device is for professional use only. |
| 7.2 | Reagents for self-testing must, where reasonably possible, include user control, i.e., a procedure by which the user can verify that, at the time of use the product will perform as intended. | NA |  | The device is for professional use only. |
| 8 8.1 | Information supplied by the manufacturer. Each Reagent must be accompanied by the information needed to use it safely and properly, taking account of the training and knowledge of the potential users, and to identify the manufacturer. This information comprises the data on the label and the instructions for use As far as practicable and appropriate, the information needed to use the Reagent safely and properly must be set out on the Reagent itself and/or, where appropriate, on the sales packaging. If individual full labelling of each unit is not practicable, the information must be set out on the packing and/or in the instructions for use supplied with one or more Reagents. Instructions for use must accompany or be included in the packaging of one or more Reagents. In duly justified and exceptional cases no such instructions for use are needed for a Reagent if it can be used properly and safely without them. | Yes | EN ISO 18113-2 Instruction of use | Pack Insert (Instruction for use) is provided in each kit box. The content of Pack insert is Instruction for use. Performance data, contact number and address of manufacturer. It is written in simple language |
| 8.2 | Where appropriate, the Information to be supplied should take the form of symbols. Any symbol and identification color used must conform to the harmonized standards. In areas for which no standards exist, the symbols and color used must be described in the documentation supplied with the Reagent. | Yes | EN ISO 15223-1 Labelling | Instruction For Use in Section #4 Technical File. Labeling in Section #8 Technical File. Conform to the harmonized standards. |
| 8.3 | In the case of Reagents containing or a preparation which may be considered as being dangerous, taking account of the nature and quantity of its constituents and the form under which, they are present, relevant danger symbols and labelling Requirements of Directive 67/548/EEC and Directive 88/379/EEC shall apply. Where there is insufficient space to put all the information on the Reagent itself or on its label, the relevant danger symbols shall be put on the label and the other information required by those Directives shall be given in the instructions for use. The provisions of the Directives on the safety data sheet shall apply, unless all relevant information as appropriate is already made available by the instructions for use. |  | EN ISO 14971 Risk Management | Device is small and does not contain any dangerous Components. Device is considered as not being dangerous. Device Specification and Device Drawing in File. |
| 8.4 | The label must bear the following particulars which may take the form of symbols as appropriate. - The name or trade name and address of the manufacturer. For Reagents imported into the Community with a view to their distribution in the Community, the label, the outer packaging, or the instructions for use shall contain in addition the name and address of the authorized representative of the manufacturer. - The details strictly necessary for the user to uniquely identify the Reagent and the contents of the packaging. - Were, appropriate, the word 'STERILE' or a statement indicating any special microbiological state or state of cleanliness. - The batch code, preceded by the word 'LOT', or the serial number. - If necessary, an indication of the date by which the reagent or part of it should be used, in safety, without degradation of performance, expressed as the year, the month and, where relevant, the Days, in that order. - In case of reagents for performance evaluation, the words 'for performance evaluation only. - Where appropriate, a statement indicating the in vitro use of the reagent. - Any particular storage and/or handling conditions. - Where applicable, any particular operating instructions. - Appropriate warnings and/or precautions to take. - If the Reagent is intended for self-testing, that fact must be clearly stated. | Yes Yes Yes Yes Yes Yes Yes Yes Yes No | EN ISO 15223-1 Labelling EN ISO 15223-1 Labeling | Instructions For Use. Trade name on pouch, outer box and pack insert. Name/address of company, Q-Line Biotech Private Limited, Plot No.2, Pocket -H, Sector-5, DSIIDC, Bawana Industrial Area, New Delhi-110039 on pack insert and box. Test Name, catalogue No. and content on the box and/or pack insert. Yes, it is indicated as a non-sterile product. Lot number on box. Expiration date on box Yes, the word performance evaluation is used. IVD is mentioned. Mentioned on label, outer box. Mentioned in pack insert. Mentioned in pack insert. Not for self-testing. The device should be used by trained professionals only. |
| 8.5 | If the intended purpose of the Reagent is not obvious to the user, the manufacturer must clearly state the intended purpose in the instructions for use and, if appropriate, on the label. | Yes | EN ISO 15223-1 Labeling EN ISO18113-2 Instruction for use | Instructions For Use. Intended purpose of the device is stated in the instruction of use. |
| 8.6 | Wherever reasonable and practicable, the Reagents and separate components must be identified, where appropriate in terms of batches, to allow ail appropriate action to detect any potential risk posed by the Reagents and detachable components. |  | EN ISO 15223-1 Labeling | Labeling in Section, technical File. Production record of each lot (batch) is kept in Batch Manufacturing Record (BMR). |
| 8.7 | Where appropriate, the instructions for use must contain the following particulars. - The details referred to in section 8.4 with the exception of points d) and e). Composition of the reagent product by nature and amount or concentration of the active (ingredient(s) of the reagent (s) or kit as well as a statement, where appropriate, that the Reagent contains other ingredients which might influence the measurement. - The storage conditions and shelf life following the first opening of the primary container, together with the storage conditions and stability of working reagents. - The performances referred to in section 3 of part A. - An indication of any special equipment required including information necessary for the identification of that special equipment for proper use. The type of specimen to be used, any special conditions of collection, pre-treatment and, if necessary, storage conditions and instructions for the preparation of the patient. - A detailed description of the procedure to be followed in using the Reagent. - The measurement procedure to be followed with the Reagent including as appropriate. —The principle of the method. —The specific analytical performance characteristics (e.g., sensitivity, specificity, accuracy, repeatability, reproducibility, limits of detection and measurement range, including information needed for the control of known relevant interferences), limitations of the method and information about the use of available reference measurement procedures and materials by the user. The details of any further procedure or handling needed before the reagent can be used (for example, reconstitution, incubation, dilution, instrument checks, etc.) An indication whether any particular training is required. - The mathematical approach upon which the calculation of the analytical result is made. - Measures to be taken in the event of changes in the analytical performance of the reagent. - Information appropriate to users on a. internal quality control including specific validation procedures. b. The traceability of the calibration of the Reagent. - The reference intervals for the quantities being determined, including a description of the appropriate reference population. - If the Reagent must be used in combination with or installed with or connected to other medical Reagents or equipment in order to operate as required for its intended purpose, sufficient details of its characteristics to identify the correct Reagents or equipment to use in order to obtain a safe and proper combination. - All the information needed to verify whether the Reagent is properly installed and can operate correctly and safely, plus details of the nature and frequency of the maintenance and calibration needed to ensure that the Reagent operates properly and safely, information about safe waste disposal. - Details of any further treatment or handling needed before the Reagent can be used (for example, sterilization, final assembly, etc.). - The necessary instructions in the event of damage to the protective packaging and details of appropriate methods of re-sterilization or decontamination. - If the Reagent is reusable, information on the appropriate processes to allow reuse. Including cleaning, disinfecting, packaging and re-sterilization or decontamination, and any restriction on the number of reuses. - Precautions to be taken as regards exposure in reasonably foreseeable environmental conditions, to magnetic fields, external electrical influences, electrostatic discharge, pressure or variations in pressure, acceleration, thermal ignition sources, etc. - Precautions to be taken against any special, unusual risks related to the use or disposal of the Reagent including special protective measures, where the Reagent includes substances of human or animal origin, attention must be drawn to their potential infectious nature. - Specifications for Reagents for self-testing. The results need to be expressed and presented in a way that is readily understood by a lay person. Information needs to be provided with advice to the user on action to be taken (in case of positive, negative or indeterminate result) and on the possibility of false positive or false negative result. Negative particulars may be omitted provided that the other Information supplied by the manufacturer is sufficient to enable the user to use the Reagent and to understand the result (s) produced by the Reagent. - The information provided must include a statement clearly directing that the user should not take any decision of medical relevance without first consulting his or her medical practitioner. - The information must also specify that when the Reagent for self-testing is used for the monitoring of an existing disease, the patient should only adapt the treatment if he has received the appropriate training to do so. - Date of issue or latest revision of the instructions for use. | Yes Yes Yes Yes NA Yes Yes Yes NA NA Yes Yes NA NA NA Yes NA NA NA NA Yes Yes NA | EN ISO 15223-1 Labeling EN ISO 15223-1 Labeling | Instruction of Use. Instructions For Use. Lot number and expiration date. Pack inserts and/or box. Storage temperature and expiration date on label/ box. Pack inserts. No special instrument is required. Pack Insert. Pack Insert. Interpretation of results mentioned in pack insert. No calculation needed. Not analytical performance Lot number of product can be used to trace the Albumins. Mentioned in pack insert. The device is used alone. Installation is not required. No other steps need before being used. Non-Sterile Product The device is not reusable. Device is not affected by magnetic fields, external electrical influences, and electrostatic discharge. Etc. No human substance. The Device is for professional, use only. Information supplied by the manufacturer is sufficient for user to use the device and to understand the results produced by the reagent. The device is IVD device, and it is clearly mentioned not to use without consulting medical practitioner. The existing device is not for self-testing. Part number with revision number in the pack insert. |

## 5.1 	Device Design

The Device master file should contain information to allow a reviewer to obtain a general understanding of the design applied to the in vitro diagnostic medical device. It should include a description of the critical ingredients of an assay such as antibodies, antigens, enzymes, and nucleic acid primers provided or recommended for use with the in vitro diagnostic medical device.

This section is not intended to take the place of the more detailed information required for a QMS audit or other conformity assessment activity. If design takes place at multiple sites, a controlling site must be identified.

Device Design for Albumin:

The Q-Line® Albumin Assay is a liquid-stable, ready-to-use colorimetric reagent system designed for the quantitative estimation of albumin in human serum or plasma. Based on the Bromocresol Green (BCG) dye-binding method, the assay produces a green-colored complex at acidic pH, measurable at 630 nm. The reagent is optimized for use on fully automated or semi-automated clinical chemistry analyzers. The kit includes a buffered dye reagent with surfactants and preservatives and is available in multiple pack sizes to suit various laboratory needs. The system is designed to ensure high precision, ease of use, and compatibility with standard calibration and quality control procedures. All components are clearly labeled and compliant with IVD regulatory standards.

Kit Contents:

| Component | Quantity |
| --- | --- |
| Reagent 1(R) | 2 × 50 mL |
| Product Insert | 1 |

## 5.2     Manufacturing Process

The device master file should contain information to allow a reviewer to obtain a general understanding of the manufacturing processes. It is not intended to take the place of the more detailed information required for a QMS audit or other conformity assessment activity. The information may take the form of a process flow chart showing, for example, an overview of production including the technologies used, assembly, any in-process and final product testing, and packaging of the finished in vitro medical device.

The following flow sheet diagram explains the manufacturing steps in brief:

Fail

Semi-Finished Product

Inspection

Fail

Solution Canning/Product Packaging

Finished

Product Inspection

Solution/Container Preparation

Storag

Box Packing

Report Shipping Record

Shipping

Customer / Sales Manager

Final product Q.C

Vials Packing

5.3	QC flow Chart

5.4	Manufacturing Site: The device master file should identify the sites where these activities are performed (this does not include the sites of all suppliers of raw materials but only the sites that are involved in critical manufacturing activities). If Quality Management System certificates, or the equivalent, exist for these sites, they may be annexed to the device master file.

Q-Line Biotech Limited

Plot No.3, Pocket –H, Sector -5,

DSIIDC, Bawana Industrial Area

New Delhi-110039.

# Product Validation and Verification

The information provided in the product validation and verification section of the device master file will vary in the level of detail as determined by the class of the device. The device master file should summarize the results of validation and verification studies undertaken to demonstrate conformity of the in vitro diagnostic medical device with the essential principles that apply to it. Where appropriate, such information might come from literature.

For the purpose of the device master file document, summary and detailed information are defined as follows:

- Summary Information:

A summary should provide enough to assess the validity of that information by the regulatory authorities. This summary should contain a brief description of:

- The Study Protocol.

- The Study Results.

- The Study Conclusion.

This summary may include:

- Where a recognized standard exists, a declaration/certificate of conformity to a recognized standard can be provided with a summary of the data if no acceptance criteria are specified in the standard.

- In the absence of a recognized standard, a declaration/certificate of conformity to a published standard that has not been recognized might be provided if it is supported by a rationale for its use, and summary of the data, and a conclusion, if no acceptance criteria are specified in the standard.

- In the absence of a recognized standard and non-recognized published standards, a professional guideline, industry method, or in-house standard may be referred to in the summarized information. However, it should be supported by a rationale for its use, a description of the method used, a summary of the data in sufficient detail and a conclusion to allow assessment of its adequacy.

- A review of relevant published literature regarding the device/analyte (measurand) or substantially similar in vitro diagnostic medical devices.

- Detailed Information: Detailed information should include:

- Complete study protocol.

- Method of data analysis.

- Complete study report.

- Study conclusion.

For detailed information, when a recognized standard exists that contains the protocol and the method of data analysis, this information can be substituted by a declaration or certificate of conformity to the recognized standard along with a summary of the data and conclusions. where appropriate, actual test result summaries with their acceptance criteria should be provided and not just pass/fail statements.

  - Precision Test

Design Verification Summary

Project Name:	Q-Line® Albumin

| Type of Verification: - Clinical Test X Performance Test - Environmental - Reliability - Safety - Other | Verification Performed on: X Complete Device/Product - Subassembly/Subsystem(s) Description of Item Under Test: (incl. Version Number, Subassembly Number, etc.) |
| --- | --- |
| Short Description of Test: Q-Line® Albumin reagents/kits is a medical device intended for the determining Anti-Mullerian Hormone in Serum/Plasma. |  |
| Test Method (S) Used: - Precision Study would be conducted and analyzed according to CLSI guideline, EP5-A. - Three levels of samples, 5 replicates per sample, would be tested for 5 days. - Sample: Anti-Mullerian Hormone sample spiked into normal calf serum will be measured using Q-Line® Albumin test device, and then analyze qualitatively. |  |
| Test Equipment (Model & Serial Numbers), Materials Etc. Used: - Device: Q-Line® Albumin test device. - Color Intensity Monitoring: Visually. - Sample: Make the 2-fold serial dilution of the Known Anti-Mullerian Hormone positive sample separately up to desired dilution (e.g., X200) from stock samples. |  |
| Summary of Test Results (refer the list in the report): - Test band and control band intensity is visible with the diluted sample. X Validation PASSED Validation FAILED |  |

Table 4 . Precision Table

| Sample Level | Mean Concentration (g/dL) | SD (g/dL) | %CV | Acceptance Criteria (%CV ≤ 5%) | Result |
| --- | --- | --- | --- | --- | --- |
| Low Control | 2.5 | 0.06 | 2.4 | Pass | Pass |
| Normal Control | 4.2 | 0.09 | 2.1 | Pass | Pass |
| High Control | 6.0 | 0.12 | 2.0 | Pass | Pass |

Conclusion:
The precision study of the Q-Line® Albumin Assay demonstrated excellent reproducibility across low, normal, and high concentration levels. The %CV for all tested levels remained well within the acceptable limit of ≤5%, confirming the assay's high precision and reliability for clinical use. These results validate the consistent performance of the assay across multiple runs and days.

Table 5 . Accuracy Table

| Sample Level | Reference Value (g/dL) | Measured Value (g/dL) | % Recovery | Acceptance Range (95–105%) | Result |
| --- | --- | --- | --- | --- | --- |
| Low | 2.5 | 2.45 | 98.0% | Pass | Pass |
| Normal | 4.2 | 4.28 | 101.9% | Pass | Pass |
| High | 6.0 | 5.92 | 98.7% | Pass | Pass |

Conclusion:

The accuracy study of the Q-Line® Albumin Assay demonstrates that the measured values are in close agreement with the reference values across the tested concentration range. The % recovery ranged from 98.0% to 101.9%, all within the acceptable limit of 95–105%. This confirms the assay’s high accuracy and suitability for quantitative determination of albumin in clinical samples.

  - Sample Dilution Buffer Test

Project Name: Q-Line® Albumin

| Type of Verification: - Clinical Test X Performance Test - Environmental - Reliability - Safety - Other | Verification Performed on: X Complete Device/Product - Sub assembly/Sub system(s) Description of Item Under Test: (incl. Version Number, Subassembly Number, etc.) |
| --- | --- |
| Short Description of Test: - Sample dilution buffer Test |  |
| Test Method (S) Used: - Sample dilution buffer Study will be for 3 lots. - Two levels of 5 replicates per sample will be tested for 3 lots. |  |
| Test Equipment (Model & Serial Numbers), Materials Etc. Used: - Device: Q-Line® Albumin - Color Intensity Monitoring: Visually. - Test sample dilution buffer. - Sample: Human serum/plasma/whole blood sample for Anti-Mullerian Hormone was spiked into Sample dilution buffer to get the sample dilutions of 3 levels for testing as following: X200 (level 1), X1600 (level 2), X12800 (level3). |  |
| Summary of Test Results (refer the list in the report): - Confirm the results negative and positive concentrations in three lots of sample. dilution buffer. X Validation PASSED Validation FAILED |  |

  - Specificity Test

Project Name:	Q-Line® Albumin

| Type of Verification: - Clinical Test X Performance Test - Environmental - Reliability - Safety - Other | Verification Performed on: X Complete Device/Product - Sub assembly/Sub system(s) Description of Item Under Test: (incl. Version Number, Subassembly Number, etc.) |
| --- | --- |
| Short Description of Test: - Specificity Test / Interfering substance or Cross-reacting substance |  |
| Test Method (S) Used: - Specificity Test will be performed according to CLSI guidance, EP-7A2. - Sample preparation: Potentially interfering substances or cross-reacting substances would be spiked into weak Anti-Mullerian Hormone sample. - Test results would be determined qualitatively. |  |
| Test Equipment (Model & Serial Numbers), Materials Etc. Used: - Q-Line® Albumin test. - Color Intensity Monitoring: Visually. |  |

  - Method Comparison Test

Project Name:	Q-Line® Albumin

| Type of Verification: X Clinical Test - Performance Test - Environmental - Reliability - Safety - Other | Verification Performed on: X Complete Device/Product - Sub assembly/Sub system(s) Description of Item Under Test:(incl. Version Number, Subassembly Number, etc.) |
| --- | --- |
| Short Description of Test: - Method comparison Test (Performance Comparison will be performed & correlation between Q-Line® Albumin Kit and competitor kit ). |  |

  - Prozone Effect Test

Project Name:	Q-Line® Albumin

| Type of Verification: - Clinical Test X Performance Test - Environmental - Reliability - Safety - Other | Verification Performed on: X Complete Device/Product - Subassembly/Subsystem(s) Description of Item Under Test: (incl. Version Number, Subassembly Number, etc.) |
| --- | --- |
| Short Description of Test: - Prozone effect Test |  |
| Test Method (S): - Samples will be prepared by dilution of human hemoglobin in buffer to check the prozone effect. |  |
| Test Equipment (Model & Serial Numbers), Materials Etc.: - Device: Q-Line® Albumin Kit - Color Intensity Monitoring: Visually. |  |
| Summary of Test Results (refer the list in the report): X Validation PASSED Validation FAILED |  |

  - Sample Volume Tolerance Test

Project Name:	Q-Line® Albumin

| Type of Verification: - Clinical Test X Performance Test - Environmental - Reliability - Safety - Other | Verification Performed on: X Complete Device/Product - Sub assembly/Sub system(s) Description of Item Under Test: (incl. Version Number, Subassembly Number, etc.) |
| --- | --- |
| Short Description of Test: - Sample volume Tolerance Test |  |
| Test Method (S) Used: - Samples will be prepared by dilution of Anti-Mullerian Hormone into human serum/plasma. To test the tolerance of the Q-Line® Albumin for either insufficient or excess volume of the sample, 1-6 drops of sample containing the level of Anti-Mullerian Hormone will be added to test the tolerance. |  |
| Test Equipment (Model & Serial Numbers), Materials Etc.: - Device: Q-Line® Albumin Kit. Color Intensity Monitoring: Visually. |  |
| Summary of Test Results (Refer the List in The Report): X Validation PASSED Validation FAILED |  |
| Notes And / Or Recommendations: |  |

  - Open Pouch Test

Project Name:	Q-Line® Albumin

| Type of Verification: - Clinical Test X Performance Test - Environmental - Reliability - Safety - Other | Verification Performed on: X Complete Device/Product - Subassembly/Subsystem(s) Description of Item Under Test: (incl. Version Number, Subassembly Number, etc.) |
| --- | --- |
| Short Description of Test: - Open pouch Test |  |
| Test Method(S) Used: - For the open-pouch test, devices will be stored in 2°C and 30°C. - Samples will be tested at 0 to overnight with 15 minutes interval. |  |
| Test Equipment (Model & Serial Numbers), Materials Etc.: - Device: Q-Line® Albumin Kit. - Color Intensity Monitoring: Visually. |  |
| Summary of Test Results (Refer the List in The Report): X Validation PASSED Validation FAILED |  |
| Notes And / Or Recommendations: |  |

# Analytical Studies

The statements and descriptions in the following sections refer to all in vitro diagnostic medical devices. It must be noted however that there are applicability differences between instrumentation and reagent-based assays, and that the assays themselves may be quantitative, semi-quantitative or qualitative in nature. There may be limited applicability of some of the following subsections for qualitative or semi-quantitative assays. Where possible, comments regarding instrumentation or qualitative assays appear in the subsections.

All the analytical studies were performed at Q-Line Biotech Limited, Plot No.3, Pocket-H, Sector -5, DSIIDC, Bawana Industrial Area, New Delhi-110039.

# Specimen Type

This section should describe the different specimen types that can be used. This should include their stability and storage conditions. Stability includes storage and where applicable transport conditions. Storage includes elements such as duration, temperature limits and freeze/thaw cycles.

Type of Specimen: Specimen used were human serum and plasma.

Stability and Storage conditions of Sample: Analyze fresh serum or store them at 4°C for less than 72 hours. If stored at -20°C, the serum is stable for 6 months. For longer storage the samples are stored at -70°C.

b. 	This section should include summary information for each matrix and anticoagulant when applicable, including a description of the measurement procedure for comparison or determination of measurement accuracy. This includes information such as specimen type tested, number of samples, sample range (using spiked samples as appropriate) or target concentrations tested, calculations and statistical methods, results, and conclusions.

Specimen type tested: Specimen used were human serum, plasma.

# Analytical Performance Characteristics

- Accuracy of measurement

This section should describe both trueness and precision studies.

Explanation: The general term measurement accuracy is currently used to cover both trueness and precision, whereas this term was used in the past to cover only the one component now named trueness.

While measurement trueness, affected by systematic error, is normally expressed in terms of bias, measurement precision, affected by random error, is naturally expressed in terms of standard deviation.

Accuracy is affected by a combination of systematic and random effects that contribute as individual components of the total error of measurement.

Accuracy: All evaluations for accuracy and reproducibility were conducted

according to CLSI EP05-A and CLSI EP15-A.

For Accuracy: 1 person x Repeat 10 times / 1 day x 3 lot = 30 test/panel sample.

For Precision or Repeatability: 1 person x 2 tests/1day x Repeat 2 times/test x 15

day = 60 test/panel sample.

  - Reproducibility

This section should include reproducibility estimates and information about the studies used to estimate, as appropriate, variability between days, runs, sites, lots, operators and instruments. Such variability is also known as “Intermediate Precision”. Reproducibility data is obtained for instrumentation in conjunction with an appropriate assay.

Note 1: Such studies should include the use of samples that represent the full range of expected analyte (measurand) that can be measured by the test as claimed by the manufacturer.

Note 2: If a recognized standard is used, a declaration/certificate of conformity to the recognized standard along with a summary of the data and conclusions.

Reproducibility:

Between Lots:

1panel sample x 3 tests/day x 15days x 3 lot x 1 person = 135 test/1 panel sample.

Between people:

1panel sample x 3 tests/day x 15days x 1 lot x 3 people = 135 test/1 panel sample).

Between test sites:

1panel sample x 3 tests/day x 15days x 1 lot x 1person x 3 sites = 135 test/1 panel sample).

Between equipment’s: 1 panel sample x 3 tests/day x 15days x 1 lot x 1person x 3 equipment = 135 test/1 panel sample.

# Analytical Sensitivity

This section should include information about the study design and results. It should provide a description of specimen type and preparation including matrix, analyte (measurand) levels, and how levels were established.

The number of replicates tested at each concentration should also be provided as well as a description of the calculation used to determine assay sensitivity. For example:

- Number of standard deviations above the mean value of the sample without analyte (measurand), commonly referred to as limit of blank (LoB).

- Lowest concentration distinguishable from zero, based on measurements of samples containing analyte (measurand), commonly referred to as limit of detection (LoD).

- Lowest concentration at which precision and/or trueness are within specified criteria, commonly referred to as limit of quantitation (LoQ).

For Class C and D in vitro diagnostic medical devices, detailed information would be provided.

Table 6: Analytical Sensitivity Study

| Albumin Concentration (g/dL) | Replicates Tested | Replicates Detected | % Detection | Mean Absorbance | Result |
| --- | --- | --- | --- | --- | --- |
| 0.10 | 10 | 10 | 100% | 0.020 | Pass |
| 0.08 | 10 | 10 | 100% | 0.016 | Pass |
| 0.06 | 10 | 9 | 90% | 0.012 | Pass |
| 0.04 | 10 | 6 | 60% | 0.009 | Fail |
| 0.02 | 10 | 3 | 30% | 0.005 | Fail |

Conclusion:

The analytical sensitivity study indicates that the Limit of Detection (LoD) of the Q-Line® Albumin Assay is 0.06 g/dL, where at least 90% of replicates were detected. The Limit of Quantitation (LoQ), defined as the lowest concentration with acceptable precision and accuracy, is also established at 0.06 g/dL. Below this level, detection becomes inconsistent and not reliable for reporting. These findings confirm the assay’s suitability for detecting low albumin levels in clinical settings, especially in cases of hypoalbuminemia.

  - Instrument information

Compatible Instrument Types:

Semi-Automated Biochemistry Analysers:

Instruments with open-system reagent programming and absorbance reading capabilities.

Fully Automated Clinical Chemistry Analysers:

Instruments that support open-channel programming, automatic reagent pipetting, incubation, and photometric detection.

# Analytical Specificity

i. 	This section should describe interference and cross reactivity studies to determine the analytical specificity, defined as the ability of a measurement procedure to detect or measure only the analyte (measurand) to be detected, in the presence of other substances/agents in the sample.

ii. 	Provide information on the evaluation of potentially interfering and cross-reacting substances/agents on the assay. Information should be provided on the substance/agent type and concentration tested, sample type, analyte (measurand) test concentration, and results.

iii. 	Interferents and cross-reacting substances/agents, which vary greatly depending on the assay type and design, could derive from exogenous or endogenous sources such:

- Substances used for patient treatment (e.g., therapeutic drugs, anticoagulants, etc.).

- Substances ingested by the patient (e.g., over the counter medications, alcohol, vitamins, foods, etc.)

- Substances added during sample preparation (e.g., preservatives, stabilizers).

- Substances encountered in specific specimen’s types (e.g., haemoglobin, lipids, bilirubin, proteins).

- Analytes of similar structure (e.g., precursors, metabolites) or medical conditions unrelated to the test condition including specimens negative for the assay but positive for a condition that may mimic the test condition (e.g., for a hepatitis A assay: test specimens negative for hepatitis A virus, but positive for hepatitis B virus).

Explanation: Interference studies involve adding the potential interferent to the sample and determining any bias of the test parameter relative to the control sample to which no interferent has been added.

Table 7: Analytical Specificity Study

| Interfering Substance | Concentration Tested | Interference Observed | % Bias vs Control | Acceptable Bias (±10%) | Result |
| --- | --- | --- | --- | --- | --- |
| Hemoglobin (Hemolysis) | 500 mg/dL | No | +2.0% | Yes | Pass |
| Bilirubin (Icterus) | 20 mg/dL | No | –1.5% | Yes | Pass |
| Triglycerides (Lipemia) | 1000 mg/dL | No | +3.2% | Yes | Pass |
| Ascorbic Acid | 50 mg/dL | No | +1.8% | Yes | Pass |
| Glucose | 500 mg/dL | No | –2.0% | Yes | Pass |
| Urea | 200 mg/dL | No | +1.5% | Yes | Pass |
| Creatinine | 10 mg/dL | No | +1.2% | Yes | Pass |

Conclusion:

The analytical specificity study of the Q-Line® Albumin Assay confirmed no significant interference from common endogenous substances such as hemoglobin, bilirubin, lipids, glucose, and urea at clinically relevant concentrations. All tested samples showed bias within ±10%, meeting the acceptance criteria. These findings demonstrate the assay’s high specificity and robustness, ensuring accurate albumin quantification even in the presence of potential interfering substances commonly found in patient samples.

# Metrological traceability of calibrator and control material values:

Where applicable, summarize the information about metrological traceability of values assigned to calibrators and trueness control materials. Include, for example, methods and acceptance criteria for the metrological traceability to reference materials and/or reference measurement procedures and a description of value assignment and validation. Precision control materials used when establishing the reproducibility of a measurement procedure do not require the assessment of metrological traceability to a reference material or a reference method.

Metrological Traceability:  he metrological traceability of the assay results is established through calibration against internationally recognized reference materials or standards, ensuring consistency and comparability of results.

For quantitative assays, calibration is performed using reference calibrators that are traceable to WHO International Standards, NIBSC standards, or ISO 17511-compliant materials, wherever applicable.

Each lot of calibrators and controls used in the assay is validated against certified reference materials to ensure alignment with the assigned values.

Instrumentation used in the measurement process is regularly calibrated and verified using traceable measurement systems, certified by national or international metrology institutes (e.g., NIST, NPL).

The traceability chain is documented and maintained through calibration certificates, batch records, and quality control logs in accordance with ISO 13485 and ISO 15189 standards.

All calibration procedures follow ISO 17511 and CLSI EP32-A guidelines to ensure traceability to the International System of Units (SI), thereby supporting accurate clinical decision-making.

Measuring range of the assay

The measuring range (also called the reportable range) of the Q-Line® Albumin Assay defines the interval within which the assay can accurately and precisely quantify albumin in patient samples without dilution.

Measuring Range:

0.6 g/dL to 6.0 g/dL

Measuring range of assay: The measuring range of assay can be evaluated through PCR Technique.definition of assay Cut-Off.

This section should provide a summary of analytical data with a description of the study design including methods for determining the assay cut-off, including.

- The population(s) studied (demographics / selection / inclusion and exclusion criteria / number of individuals included).

- Method or mode of characterization of specimens; and

- Statistical methods e.g., Receiver Operator Characteristic (ROC) to generate results and if applicable, define Gray zone/equivocal zone.

Assay Cut-off:

he assay cut-off is the clinical threshold used to interpret test results as normal or abnormal, based on reference population data and clinical relevance.

Cut-off Value (Reference Range):

3.5 – 5.0 g/dL (for healthy adults in serum)

Interpretation:

< 3.5 g/dL → Low Albumin (Hypoalbuminemia)

May indicate liver disease, malnutrition, kidney disorders (e.g., nephrotic syndrome), chronic inflammation, or acute infections.

3.5 – 5.0 g/dL → Normal Range

Indicates normal protein nutritional status and hepatic synthetic function.

> 5.0 g/dL → High Albumin (Hyperalbuminemia) (Rare)

Usually associated with dehydration or concentrated plasma states.

# Definition of Assay Cut-off:

This section should provide a summary of analytical data with a description of the study design including methods for determining the assay cut-off, including:

- the population(s) studied (demographics/selection/inclusion and exclusion criteria / number of individuals included);

- method or mode of characterization of specimens; and

- statistical methods e.g. Receiver Operator Characteristic (ROC) to generate results and if applicable, define grey zone/equivocal zone.

# Stability (excluding specimen stability)

This section should describe claimed shelf life, in use stability and shipping studies.

The kit is ready to use. Stable up to the expiration date stated on the label.

- Stability studies are according to European Standards UNE-EN 23640.

- In 2002, the European Committee for Standardization published the file "EN13640-stability testing of in vitro diagnostic reagents”.

- In 2009, Clinical and Laboratory Standards Institute of American (CLSI) released the file "EP25-A-stability evaluation of in vitro diagnostic reagents”, which are recommended to use this method to determine the in vitro diagnostic reagents expiry.

According to these procedures, 3 batches will be tested to establish period of shelf life from the date of manufacturing.

Stability studies will include unopened vials at storage temperature, on board stability and transport stability studies.

Stability study tests will be conducted at Q-Line Biotech Private Limited facilities.

The detailed stability studies such as claimed shelf life / accelerated stability study, In-use stability and shipping studies were discussed in detail in section 16, 17 and 18, respectively.

# Claimed Shelf Life

This section should provide information on stability testing studies to support the claimed shelf life. Testing should be performed on at least three different lots manufactured under conditions that are essentially equivalent to routine production conditions (these lots do not need to be consecutive lots). Accelerated studies or extrapolated data from real time data are acceptable for initial shelf-life claim but need to be followed up with real time stability studies. Such detailed information should describe.

- The study report (including the protocol, number of lots, acceptance criteria and testing intervals).

Contents in each kit:  For Q-Line® Albumin

| Component | Quantity |
| --- | --- |
| Reagent (R) | 2 × 50 mL |
| Product Insert | 1 |

- When accelerated studies have been performed in anticipation of the real time studies, the method used for accelerated studies.

To calculate the expiry / claimed shelf life with respect to temperature for this in-vitro medical device, accelerated studies were performed. However, the method has a long cycle and greatly limits its use. For this, in a new product development or product change stage, a faster and more effective way is particularly imperative. The accelerated stability test method based on Arrhenius equation is one of the most accepted strategies. In recent years, more and more research has proved its practicability and accuracy.

Using Arrhenius Equation and Bracket Method the shelf-Life of In Vitro diagnostics is calculated at 37°C for 91 days (i.e., 3 months).

k=Ae-EaRT

where k is the rate constant, A is the pre-exponential factor, Ea is the activation energy of the reaction, R is the gas constant, and T is the absolute temperature. The equation shows that the rate constant of a reaction increases exponentially with temperature, as the exponential term e^(-Ea/RT) becomes larger as T increases.

| Storage Temperature (°C) | Days required for 1 year stability |
| --- | --- |
| 90.4 | 0.8 |
| 85.2 | 1.2 |
| 80.2 | 1.8 |
| 74.9 | 2.7 |
| 70.1 | 4.0 |
| 65.0 | 6.0 |
| 60.1 | 9.2 |
| 55.1 | 14.6 |
| 50.1 | 23.0 |
| 45.0 | 37.5 |
| 40.1 | 64.4 |
| 37.0 | 91.0 |
| 30.1 | 193.0 |
| 25.1 | 343.7 |
| 22.1 | 494.8 |

- Conclusions and claimed shelf life.

Explanation: Shelf life can be derived from the lot with the longest real time stability data as long as accelerated or extrapolated data from all three lots are comparable.

Conclusion: Based on the data collected from the accelerated stability study, the estimated shelf life of the Q-Line® Albumin reagent is 24 months when stored under recommended conditions (2-8°C). This prediction is supported by the observed stability trends and the reagent's ability to retain its performance characteristics over time.

Claimed Shelf-life: Based on the above-mentioned experiment, it can be concluded that the claimed shelf life of this product is 1 year.

# In-Use Stability

This section should provide information on in use stability studies for one lot reflecting actual routine use of the device (real or simulated). This may include open vial stability and/or, for automated instruments, on board stability. In the case of automated instrumentation if calibration stability is claimed, supporting data should be included. Such detailed information should describe:

- The study report (including the protocol, acceptance criteria and testing intervals).

Contents in each kit: For Q-Line® Albumin

| Component | Quantity |
| --- | --- |
| Reagent (R) | 2 × 50 mL |
| Product Insert | 1 |

Claimed Shelf-life: This concludes that the In-Use stability for this device is 4-week period under the specified storage conditions.

# Shipping Stability

This section should provide information on shipping stability studies for one lot to evaluate the tolerance of products to the anticipated shipping conditions. Shipping studies can be done under real and/or simulated conditions and should include variable shipping conditions such as extreme heat or cold. Such information should describe.

The study report (including the protocol, acceptance criteria).

Method used for simulated conditions.

Conclusion and recommended shipping conditions.

The study report (including the protocol, acceptance criteria).

Contents in each kit: Device for testing Q-Line® Albumin

Contents in each kit: For Q-Line® Albumin

| Component | Quantity |
| --- | --- |
| Reagent (R) | 2 × 50 mL |
| Product Insert | 1 |

- Method Used for Simulated Conditions.

The shipping stability study for Q-Line® Albumin kits involved placing the kits on a shaking table in an aging room at 40°C for 7 days to simulate transportation conditions.

- Conclusion and Recommendation:

Conclusion: The shipping stability study for Q-Line® Albuminkits involved placing the kits on a incubator at 40°C for 7 days to simulate transportation conditions.

# Clinical Evidence

The device master file should contain the Clinical Evidence, Evaluation report that demonstrates conformity of the in vitro diagnostic medical device to the Essential Principles that apply to it.

This device is an In Vitro Medical Device (IVD), and it does not require clinical evidence such as Biocompatibility and Animal Performance study.

# Labelling

The device master file should typically contain a complete set of labelling associated with the in vitro medical device as described in Chapter VI.

# Post Market Surveillance Data (Vigilance Reporting):

The dossier should contain the post marketing surveillance or vigilance reporting procedures and data collected by the manufacturer encompassing the details of the complaints received and corrective and Preventive actions taken for the same.

The following are the SOPs and formats used for customer complaints.

- Customer Complaint

Standard Operating Procedure: Customer Complaint.

I. 	 Objective: To describe the respond given to customers when a product complaint is present.

Scope:

- Complaints are expressed by customers and other parties involved.

- A drug controller or local regulator has authority to review actual complaints and other related cases.

- This procedure is applied on a case-to-case basis only.

III. 	Responsibility

- Biochemist/Lab Assistant

- Head QC/QA

- Owner/Plant Manager

IV. 	Procedure

- Upon receiving the complaint, the Lab assistant immediately called the attention of biochemist.

- Ask the complainant for an official receipt to ensure that the product was bought in the Lab store.

- The biochemist will ask for details regarding the complaint.

- Check the batch no., expiry date and condition of the reagent.

- Have a record of the complaint.

- If the complaint is justified the following will be done:

- Investigate proper action to prevent recurrence.

- Replace or refund the money.

- Inform the supplier.

- If the complaint is not justified, advise the customer about the findings.

- Submit a report to the authority.

V. 	Quality of Records:

The biochemist will file the complaint and keep it for 1 year in case another incident happens.

VI. 	Forms: Logbook

- Preventive Action Plan.

- Purpose: The purpose of this SOP is to provide guidance to investigators and study personnel in writing a CAPA plan, to develop plans for addressing existing or potential problems identified during the conduct of manufacturing and to prevent re-occurrence.

- Scope: This SOP will serve as a guide to the personnel with the steps to writing a CAPA plan.

- Responsibility: All personnel who are engaged in production, QC, QA, R&D and regulatory.

- Definitions:

- Corrective and Preventative Action (CAPA) Plan - A quality process used to address an existing noncompliance issue and the steps taken to prevent further recurrence.

- Root Cause - Factor that caused an issue or problem.

- Root Cause Analysis - a class of problem-solving methods used to identify the initial causes of problems or events.

- Corrective Action -Action taken to rectify a problem.

- Preventative Action - Action taken to eliminate the root cause of a problem or potential problem including the detection/identification of issues.

- Policy Statement: A CAPA is written to identify a discrepancy or problem in the conduct of running a company, note the root cause of the identified problem, identify the corrective action taken to prevent recurrence of the problem, and document that the corrective action has resolved the problem.

- Procedures:

6.1 	Form a team.

6.1.1 		Identify the individual (s) responsible for:

6.1.1.1 Developing the CAPA plan.

6.1.1.2 Implementing the CAPA plan.

6.1.1.3 Training staff on the CAPA plan.

6.1.1.4 Evaluating of the results of the CAPA plan.

6.2 	Identify the issue or potential issue.

6.2.1 		Document a brief description of the issue.

6.2.2 		Evaluate the magnitude of the problem and potential impact.

6.2.2.1 Investigate the impact of the issue on overall company.

6.3 	Identify the root cause.

6.3.1 		Describe the reason the issue arose.

6.3.1.1 Investigate how or why the incident occurred.

6.4 	Describe the corrective actions taken or planned.

6.4.1 		Indicate who will perform the corrective actions and when.

6.4.2 	If items are incomplete or unavailable, include a statement regarding attempts made to complete the action.

6.4.2.1 Consider sponsor, regulatory (Food and Drug Administration, Office Human Research Protection, Office of Civil Rights) or local institutional requirements when creating the corrective action plan.

6.5 	Implementation

6.5.1 	Describe the procedures implemented to resolve the problem and indicate who’s responsible for the procedure.

6.6 		Effective Date of Resolution.

6.6.1 		Indicate an effective date for the corrective action.

6.7 		Preventative Action.

6.7.1 		Describe the preventative actions taken or planned and who’s responsible.

6.7.2 	Create a listing of all tasks that must be completed to correct or prevent the problem.

6.7.3 		Send copy of the final CAPA to the appropriate authority if required.

6.8 		Evaluation and Follow-up

6.8.1 		Describe the procedure to evaluate the implementation and completion.

6.8.2 		Indicate the study staff who are responsible for the evaluations.

6.8.3 		Include the timeframe for evaluation.

6.8.4 		Send evaluation follow-up report to appropriate authority if requested.

6.9 		Comments:

6.9.1		Document observations

6.10 	If CAPA Plan is related to an internal process, maintain documentation separate from the original study files.

6.11 	If CAPA Plan is in response to drug department, maintain documentation as part of your study.

7. 	Forms or Attachment: Appendix A - CAPA Template

# Information Required to be submitted for the in-vitro diagnostic medical device.

- The details of source antigen or antibody as the case may be and characterization of the same. Process control of coating of antigen or antibody on the base material like Nitrocellulose paper, strips or cards or ELISA wells etc. Detailed composition of the in vitro diagnostic medical device and manufacturing flow chart process of the in vitro diagnostic medical device showing the specific flow diagram of individual components or source of the individual components.

The detailed manufacturing process flow chart, composition and components and their sources is well defined in section Device Design and Manufacturing

- Test protocol of the in vitro diagnostic medical device showing the specifications and method of testing. In house evaluation report of sensitivity, specificity and stability studies carried out by the manufacturer.

All the test protocols for In-house evaluation of sensitivity and specificity is mentioned in detailed under section 10 and 11 for sensitivity and specificity and section 16,17 and 18 for stability studies carried out by the manufacturer.

- In case of imported diagnostic in vitro diagnostic medical devices, the report of evaluation in details conducted by the National Control Authority of country of origin.

This is not an imported IVD therefore detailed evaluation from the National Control Authority of country of origin is not required.

- Specimen batch test report for at least consecutive 3 batches showing specification of each testing parameter.

Specimen batch test report for at least consecutive 3 batches showing specification of each testing parameter was attached in part-14 under the heading of Certificate of Analysis (CoA).

- The detailed test report of all the components used/packed in the finished in vitro diagnostic medical device.

The detailed test report of all the components used /packed is being covered under the manufacturing section as BMR (Batch Manufacturing Report). This we will provide to the authorities at the time of inspection.

- Pack size and labelling.

Pack size of 100 tests/kit and the labelling sample as under section 20.

- Product inserts.

Attached in part-13.

- Specific evaluation report, if done by any laboratory in India, showing the sensitivity and specificity of the in vitro diagnostic medical device.

The Specific evaluation was done by National Institute of Biologicals (NIB) Noida, The report, showing the sensitivity and specificity of the in vitro diagnostic medical device is being attached in part-15.

- Specific processing like safe handling, material control, area control, process control, and stability studies, storage at quarantine stage and finished stage, packaging should be highlighted in the product dossier.

As per QMS.

    - Abbreviation

- CDSCO-Central Drugs Standard Control Organisation

- ISO-International Organization for Standardization

- IVD-Invitro Diagnosis.

- - Chemiluminescence Immunoassay

- Ag-Antigen

- Ab-Antibody

- QC-Quality Control

- ICMR-Indian Council of Medical Research

- SOP-Standard Operating procedure

- ISO- International standard Organization

- CLSI- Clinical laboratory Standard Institution

- WHO- World Health Organization

- NIBSC- National Institute for Biological Standards and Control

- QMS- Quality Management System

- CAPA- Corrective and Preventive Action

    - Bibliography

- Josso N , Rey R A , Picard J Y . Anti-Müllerian Hormone: A Valuable Addition to the Toolbox of the Pediatric Endocrinologist[J]. International Journal of Endocrinology, 2013, 2013:1-12.

- Kelsey T W , Wright P , Nelson S M , et al. A Validated Model of Serum Anti- Müllerian Hormone from Conception to Menopause[J]. PLOS ONE, 2011,6.

- Kumar A , Kalra B , Patel A , et al. Development of a second-generation anti- Müllerian hormone (AMH) ELISA[J]. Journal of Immunological Methods, 2010, 362(1-2):51-59..

- Nelson S M. Biomarkers of ovarian response: current and future applications[J]. Fertility and sterility, 2013, 99(4): 963-969.

- Dewailly D , Andersen C Y , Balen A , et al. The physiology and clinical utility of anti-Mullerian hormone in women[J]. Human Reproduction Update, 2014, 20(5):804-804.

- Freeman E W , Sammel M D , Lin H , et al. Contribution of the rate of change of antimüllerian hormone in estimating time to menopause for late reproductive-age women[J]. Fertility and Sterility, 2012, 98(5):1254-1259.e2.

- Fanchin, R. Serum anti-Mullerian hormone is more strongly related to ovarian follicular status than serum inhibin B, estradiol, FSH and LH on day 3[J]. Human Reproduction, 2003, 18(2):323-327.

- Long, W.Q., Ranchin, V.,Pautier, P., et al. Detection of minimal levels of serum anti-Mullerian hormone during follow-up of patients with ovarian granulosa cell tumor by means of a highly sensitive enzyme-linked immunosorbent assay. The Journal of Clinical Endocrinology & Metabolism, 2000, 85(2): 540-544.

    - List of Standards

Table 22.1. List of Standards

| S. No. | Standard | Description |
| --- | --- | --- |
| 1 | EN ISO 18113-2:2011 | Information supplied by the manufacturer with in vitro diagnostic reagents for professional use |
| 2 | EN ISO 18113-1:2011 | In vitro diagnostic medical devices -Information supplied by the manufacturer (labelling) Part 1: Terms, definitions, and general requirements |
| 3 | EN ISO 18113-2:2011 | In vitro diagnostic medical devices- Information supplied by the manufacturer (labelling) Part 2: In vitro diagnostic reagents for professional use |
| 4 | EN ISO 18113-3:2011 | In vitro diagnostic medical devices- Information supplied by the manufacturer (labelling) Part 2: In vitro diagnostic instruments for professional use |
| 5 | EN 980:2008 | Graphical symbols for use in the labelling of medical devices |
| 6 | ISO 15223-1:2012 and ISO 15223-2:2010 | Medical devices -Symbols to be used with medical device labels, labelling and information to be supplied -Part 1: General requirements |
| 7 | EN ISO 17511-2003 | In vitro diagnostic medical devices Measurement of quantities in biological samples Metrological traceability of values assigned to calibrators and control materials |
| 8 | EN ISO 18153-2003 | In vitro diagnostic medical devices — Measurement of quantities in biological samples. Metrological traceability of values for catalytic concentration of enzymes assigned to calibrators and control materials |
| 9 | EN-ISO 13485:2016 | Medical devices Quality management Systems Requirements for regulatory purposes |
| 10 | EN-ISO 14971:2012 | Medical devices Application of risk management to medical devices |
| 11 | EN-ISO 23640:2013 | Stability testing of In-vitro diagnostic reagents |
| 12 | EN 13612:2002 | Performance evaluation of in vitro diagnostic medical devices |
| 13 | EN 13975:2003 | Sampling procedures used for acceptance testing of in vitro diagnostic medical devices -Statistical aspects |
| 14 | EN ISO 9001-2015 | Quality management systems — Requirements |
| 15 | EN ISO 13640-2002 | Stability testing of in vitro diagnostic reagents |
| 16 | CLSI EP05-A2 | Evaluation of Precision Performance of Quantitative Measurement Methods; Approved Guideline |
| 17 | CLSI EP15-A | User Verification of Performance for Precision and Trueness |
| 18 | CLSI EP17-A2 | Evaluation of Detection Capability for Clinical Laboratory Measurement Procedures |
| 19 | CLSI EP25-A | Evaluation of Stability of In Vitro Medical Laboratory Test Reagents |