QUALITY ASSURANCE DEPARTMENT CERTIFICATE OF ANALYSIS

| Name of the Product | INDO-MEDX® Aura Cal Biochemistry Multical |  |  |
| --- | --- | --- | --- |
| Model | - | Batch Number | CAL-Bio-TB-01 |
| MFG Date | MAY-2026 | EXP Date | APR-2028 |
| Quality Release date | 07/05/2026 |  |  |

| S. No | Parameter | Method | Unit | Target | Rep 1 | Rep 2 | Rep 3 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 1 | Albumin | Bromocresol Green | g/dl | 3.08 | 3.05 | 3.11 | 3.06 |
| 2 | Alkaline Phosphatase | AMP optimised to IFCC 37°C | U/L | 361 | 361 | 361 | 360.7 |
| 3 | ALT/SGPT | IFCC without PDP 37°C | U/L | 121 | 121.7 | 121.4 | 120.1 |
| 4 | Amylase | Direct (CNPG3) | U/L | 277 | 277.7 | 276.9 | 277.6 |
| 5 | AST/SGOT | IFCC without PDP 37°C | U/L | 98 | 98.38 | 98.16 | 98.16 |
| 6 | Bilirubin direct | Diazo Method | mg/dl | 1.94 | 1.93 | 1.94 | 1.94 |
| 7 | Bilirubin total | Diazo Method | mg/dl | 4.98 | 4.98 | 4.99 | 4.97 |
| 8 | Calcium | Arsenazo III Method | mg/dl | 12.2 | 12.2 | 12.23 | 12.27 |
| 9 | Chloride | Colorimetric | Mmol/L | 111 | 110.7 | 111.7 | 111.1 |
| 10 | Cholesterol | CHOD/POD Method | mg/dl | 286 | 286.9 | 286.9 | 285.8 |
| 11 | CK-NAC | IFCC 37°C | U/L | 513 | 512.8 | 512.6 | 513.8 |
| 12 | Creatinine | Alkaline Picrate | mg/dl | 3.55 | 3.56 | 3.54 | 3.55 |
| 13 | Creatinine | Jaffe | mg/dl | 3.48 | 3.49 | 3.47 | 3.49 |
| 14 | Creatinine | Enzymatic | mg/dl | 3.71 | 3.71 | 3.73 | 3.73 |
| 15 | gamma-GT | IFCC 37°C | U/L | 173 | 171.2 | 172.7 | 172.6 |
| 16 | Glucose | GOD/POD Method | mg/dl | 277 | 277.5 | 277.1 | 279 |
| 17 | Iron | Colorimetric without ppt | µg/dl | 210 | 210.2 | 209.7 | 210.5 |
| 18 | LDH | L to P IFCC | U/L | 331 | 330.4 | 327.8 | 334.3 |
| 19 | Lipase | Colorimetric without ppt | U/L | 71 | 70.62 | 71.62 | 70.28 |
| 20 | Magnesium | Xylidyl Blue | mg/dl | 4.06 | 4.08 | 4.06 | 4.06 |
| 21 | Phosphorus | Phosphomolybdate UV | mg/dl | 7.22 | 7.29 | 7.24 | 7.21 |
| 22 | Total protein | Biuret End Point | g/dl | 4.58 | 4.57 | 4.58 | 4.58 |
| 23 | TIBC | Colorimetric | µg/dl | 255 | 255.7 | 254.4 | 254.5 |
| 24 | Triglycerides | GPO/PAP No Correction | mg/dl | 264 | 266.4 | 266.6 | 266.6 |
| 25 | Urea | GLDH Method | mg/dl | 115 | 114.4 | 115.1 | 114.5 |
| 26 | Uric Acid | AOX Method | mg/dl | 9.04 | 9.03 | 9.05 | 9.04 |

Shelf Life: 24 months from the date of manufacture.

Storage Temperature: 2 ~ 8℃.

Conclusion: Passed.

Remark: The analytical performance of INDO-MEDX® Aura Cal Biochemistry Multical (Batch No.: CAL-Bio-TB-01) complies with internal quality specifications and acceptance criteria. We certify that this lot passes and complies with above mentioned tests and specifications.

| Tested By: |  | Approved By: |  |
| --- | --- | --- | --- |
| Name | M RAVEENKUMAR | Name | GOVIND MESTRY |
| Signature |  | Signature |  |
| Date |  | Date |  |